## Split Scan - NDF
Measurements to determine NDF %T for "split scan" method - using two NDF filters
Instructions: Sphere should be empty. Make sure you use the same excitation wavelength, slit widths and other insturment settings used for your "scatter" measurements. You need a 2nd NDF that has a similar %T to the one used in "Scatter" measurements.
Measurement 1: Put the normal "scatter" NDF filter into emission path (as if you are running a "scatter" measurement). Add a second ("new") NDF to the excitation path (similar or same value to the NDF used in front of detector for scatter measurements).
Measurement 2: Remove the "scatter" NDF filter, but keep in the "new" NDF filter (to protect the detector from becoming saturated).
Measurement 3: Block the light beam using a thick black object, to measure "dark" (background) counts.
INPUTS |  |  |  | OUTPUTS
| Measurement 1 | Measurement 2 | Measurement 3
Wavelength nm | Both filters in | "Scatter" filter out | Dark counts | Filter In | Filter Out
405 | -12833.6 | 54725.2 | -1561.0393 | -11272.5607 | 56286.2393
405.1 | -9436.5887 | 73184.3 | -11077.7 | 1641.1113 | 84262
405.2 | 6854.9657 | 46072.8 | 2510.7254 | 4344.2403 | 43562.0746
405.3 | -16223.3 | 65069.6 | 1153.4209 | -17376.7209 | 63916.1791
405.4 | 5493.5136 | 41368.2 | -4952.2568 | 10445.7704 | 46320.4568
405.5 | -6041.2717 | 73910.2 | -4953.5583 | -1087.7134 | 78863.7582
405.6 | -611.1567 | 51605.4 | 6580.3656 | -7191.5223 | 45025.0344
405.7 | -1291.2438 | 65363.2 | -2921.3193 | 1630.0756 | 68284.5193
405.8 | 4139.2453 | 77716.7 | -23932.7 | 28071.9453 | 101649.4
405.9 | 1419.7176 | 70927.6 | -2914.9261 | 4334.6437 | 73842.5262
406 | 4805.6959 | 71572.1 | -2241.7766 | 7047.4726 | 73813.8766
406.1 | 2096.1815 | 73767.7 | -3606.8739 | 5703.0553 | 77374.5739
406.2 | -4677.1 | 59379.1 | 4551.1608 | -9228.2608 | 54827.9392
406.3 | -1970.0309 | 55913.2 | 5228.0008 | -7198.0317 | 50685.1992
406.4 | 10257 | 72022.1 | 1836.8384 | 8420.1616 | 70185.2616
406.5 | -17643.4 | 81685.8 | -10453.1 | -7190.3 | 92138.9
406.6 | 2798.3488 | 73283.6 | -15181.9 | 17980.2488 | 88465.5
406.7 | -2647.0466 | 73239.9 | 7975.4257 | -10622.4723 | 65264.4743
406.8 | -10109 | 61126 | -13818.4 | 3709.4 | 74944.4
406.9 | -12850.9 | 83788 | 2515.9458 | -15366.8458 | 81272.0542
407 | 748.7255 | 61235 | 475.296 | 273.4296 | 60759.704
407.1 | 2112.5586 | 79774.4 | -7668.0538 | 9780.6123 | 87442.4538
407.2 | 2788.4705 | 72346 | -4952.0511 | 7740.5216 | 77298.0511
407.3 | -4012.9315 | 73042.8 | -4957.2765 | 944.345 | 78000.0765
407.4 | 67.9886 | 56625.3 | -12404.2 | 12472.1886 | 69029.5
407.5 | -9466.6698 | 70037.3 | -9058.0366 | -408.6332 | 79095.3366
407.6 | 751.0493 | 82293 | -16559.5 | 17310.5493 | 98852.5
407.7 | -12874.4 | 89667.1 | -9729.3657 | -3145.0343 | 99396.4657
407.8 | 8249.9814 | 86384.5 | 5258.7038 | 2991.2775 | 81125.7962
407.9 | -4707.3274 | 63492.9 | 7996.8401 | -12704.1675 | 55496.0599
408 | -8797.1255 | 98174.4 | -8412.4912 | -384.6343 | 106586.8912
408.1 | -6101.256 | 82630.4 | 8029.3106 | -14130.5666 | 74601.0894
408.2 | -6074.5262 | 84404.7 | -10462.2 | 4387.6738 | 94866.9
408.3 | 8237.2975 | 81649 | -15876.6 | 24113.8975 | 97525.6
408.4 | -1977.8839 | 104123 | -4285.9943 | 2308.1103 | 108408.9943
408.5 | -12279.8 | 90434.1 | 3884.9791 | -16164.7791 | 86549.1209
408.6 | -3355.695 | 92743 | 3202.1222 | -6557.8172 | 89540.8778
408.7 | -9463.8179 | 109718 | 12072.6 | -21536.4179 | 97645.4
408.8 | 68.2336 | 114944 | -5661.5136 | 5729.7472 | 120605.5136
408.9 | -13563.2 | 121190 | 4589.2566 | -18152.4566 | 116600.7434
409 | 3477.3573 | 105176 | 3912.5956 | -435.2383 | 101263.4044
409.1 | 2112.4236 | 143629 | 5941.3724 | -3828.9488 | 137687.6276
409.2 | -613.3634 | 124512 | 3887.4606 | -4500.824 | 120624.5394
409.3 | 68.1963 | 116770 | -2936.7136 | 3004.9099 | 119706.7136
409.4 | 10982.5 | 117979 | -3628.0632 | 14610.5632 | 121607.0632
409.5 | 3477.7595 | 139230 | 16209.6 | -12731.8405 | 123020.4
409.6 | -21835.2 | 127060 | -7732.6769 | -14102.5231 | 134792.6769
409.7 | -19140.3 | 131172 | -5002.39 | -14137.91 | 136174.39
409.8 | 753.2316 | 144918 | 7350.7779 | -6597.5463 | 137567.2221
409.9 | 4188.1028 | 146862 | -5693.6451 | 9881.7479 | 152555.6451
410 | -617.0935 | 137114 | -4365.442 | 3748.3485 | 141479.442
410.1 | -1984.9076 | 161156 | 1858.4853 | -3843.3929 | 159297.5147
410.2 | 8281.5264 | 142003 | -6384.1759 | 14665.7023 | 148387.1759
410.3 | -23861.7 | 160032 | 3897.9607 | -27759.6607 | 156134.0393
410.4 | -7524.4306 | 165272 | 5277.6854 | -12802.116 | 159994.3146
410.5 | -7453.4033 | 175899 | 1158.1642 | -8611.5675 | 174740.8358
410.6 | -1292.5777 | 159456 | -2256.7441 | 964.1664 | 161712.7441
410.7 | -6066.052 | 205490 | -12521.6 | 6455.548 | 218011.6
410.8 | 8943.6717 | 210990 | -4990.8043 | 13934.4759 | 215980.8043
410.9 | -9519.6697 | 234039 | 480.0761 | -9999.7458 | 233558.9238
411 | 6927.4564 | 226893 | 7339.963 | -412.5066 | 219553.037
411.1 | 4173.4916 | 251342 | -13900.8 | 18074.2916 | 265242.8
411.2 | 3486.4037 | 327458 | -3633.9451 | 7120.3488 | 331091.9451
411.3 | 6221.4146 | 329064 | -1579.1327 | 7800.5472 | 330643.1327
411.4 | 6234.3674 | 415711 | 8029.9907 | -1795.6232 | 407681.0093
411.5 | 16503.6 | 504980 | 8723.4989 | 7780.1011 | 496256.5011
411.6 | 4177.0873 | 702627 | 1852.6807 | 2324.4065 | 700774.3193
411.7 | 14451.9 | 879924 | 14191.1 | 260.8 | 865732.9
411.8 | 14530.3 | 1314440 | -206.0422 | 14736.3422 | 1314646.0422
411.9 | -5399.7232 | 1489260 | -7762.1601 | 2362.4368 | 1497022.1601
412 | 68.1904 | 1786220 | 3898.9428 | -3830.7524 | 1782321.0572
412.1 | 16475.6 | 2633910 | 10031.5 | 6444.1 | 2623878.5
412.2 | 17862.3 | 3200370 | -4988.6944 | 22850.9944 | 3205358.6944
412.3 | 40536.4 | 4548640 | -12515.3 | 53051.7 | 4561155.3
412.4 | 23388.1 | 5219770 | -205.3516 | 23593.4516 | 5219975.3516
412.5 | 46662.4 | 7212620 | 6655.7464 | 40006.6536 | 7205964.2536
412.6 | 39199 | 8528650 | 481.2829 | 38717.7171 | 8528168.7171
412.7 | 65331 | 11000500 | -6369.7802 | 71700.7802 | 11006869.7802
412.8 | 61076.8 | 12732400 | -7046.9513 | 68123.7513 | 12739446.9513
412.9 | 79901.6 | 14633300 | -9108.1426 | 89009.7426 | 14642408.1426
413 | 119536 | 18217900 | -4310.4996 | 123846.4996 | 18222210.4995
413.1 | 139401 | 20625200 | -205.4169 | 139606.4169 | 20625405.4169
413.2 | 144491 | 24784200 | -9780.7837 | 154271.7837 | 24793980.7837
413.3 | 174682 | 27628200 | -10486.5 | 185168.5 | 27638686.5
413.4 | 198087 | 32898100 | -890.8596 | 198977.8596 | 32898990.8597
413.5 | 229794 | 34913600 | 20349.2 | 209444.8 | 34893250.8
413.6 | 247139 | 40673700 | -9835.8344 | 256974.8344 | 40683535.8344
413.7 | 285185 | 43774100 | -18168.7 | 303353.7 | 43792268.7
413.8 | 278431 | 46261100 | -206.2062 | 278637.2062 | 46261306.2062
413.9 | 303888 | 51787600 | 10760.3 | 293127.7 | 51776839.7
414 | 326731 | 54586200 | -7043.1309 | 333774.1309 | 54593243.1309
414.1 | 384857 | 59482000 | -19403.4 | 404260.4 | 59501403.4
414.2 | 345995 | 61795500 | -10480.7 | 356475.7 | 61805980.7
414.3 | 406596 | 66823200 | -20720.8 | 427316.8 | 66843920.8
414.4 | 394945 | 68645600 | 4585.5785 | 390359.4215 | 68641014.4215
414.5 | 410696 | 72518800 | -5691.717 | 416387.717 | 72524491.717
414.6 | 448108 | 73524900 | -2949.3857 | 451057.3857 | 73527849.3857
414.7 | 404931 | 75025600 | 12132.8 | 392798.2 | 75013467.2
414.8 | 461077 | 75878500 | 6658.4756 | 454418.5244 | 75871841.5244
414.9 | 414161 | 76167000 | 3907.6287 | 410253.3713 | 76163092.3713
415 | 453186 | 76313100 | -1580.9962 | 454766.9962 | 76314680.9962
415.1 | 447334 | 75955000 | 6665.4305 | 440668.5695 | 75948334.5695
415.2 | 414308 | 74364400 | 2542.4709 | 411765.5291 | 74361857.5291
415.3 | 425776 | 72539900 | -205.0553 | 425981.0553 | 72540105.0553
415.4 | 423870 | 69384100 | 10752.5 | 413117.5 | 69373347.5
415.5 | 378643 | 67629000 | -13229.4 | 391872.4 | 67642229.4
415.6 | 414711 | 65823100 | 12826.5 | 401884.5 | 65810273.5
415.7 | 374549 | 60516100 | 4586.7507 | 369962.2493 | 60511513.2493
415.8 | 365897 | 58399700 | -2255.9983 | 368152.9983 | 58401955.9983
415.9 | 324134 | 53336000 | -4314.2989 | 328448.2989 | 53340314.2988
416 | 297998 | 50864700 | 1849.8332 | 296148.1668 | 50862850.1668
416.1 | 260306 | 45017300 | 8695.2628 | 251610.7372 | 45008604.7372
416.2 | 254157 | 41851500 | 2533.2526 | 251623.7474 | 41848966.7474
416.3 | 206849 | 36466400 | 5270.1328 | 201578.8672 | 36461129.8672
416.4 | 241019 | 33685500 | 14178.9 | 226840.1 | 33671321.1
416.5 | 196473 | 30975000 | -10521.9 | 206994.9 | 30985521.9
416.6 | 154727 | 26188400 | -8468.2746 | 163195.2746 | 26196868.2746
416.7 | 144510 | 23639600 | -13951.9 | 158461.9 | 23653551.9
416.8 | 104965 | 19178200 | 8027.3169 | 96937.6831 | 19170172.6831
416.9 | 84061.8 | 17001900 | -8392.7164 | 92454.5164 | 17010292.7164
417 | 86207.1 | 13500700 | -8389.3525 | 94596.4525 | 13509089.3525
417.1 | 76905 | 12110500 | -4302.9233 | 81207.9233 | 12114802.9233
417.2 | 67300.1 | 9114090 | -4991.5059 | 72291.606 | 9119081.506
417.3 | 56629.9 | 7848050 | -889.953 | 57519.853 | 7848939.953
417.4 | 34877.9 | 6834330 | 8686.0617 | 26191.8383 | 6825643.9383
417.5 | 11672.2 | 4996120 | -13898.7 | 25570.9 | 5010018.7
417.6 | 33541.3 | 4306160 | -8430.2411 | 41971.5411 | 4314590.2411
417.7 | 15822.9 | 3002390 | -7079.6108 | 22902.5108 | 3009469.6108
417.8 | 34941.8 | 2508090 | 14925.1 | 20016.7 | 2493164.9
417.9 | -5410.5178 | 1756670 | -10492.5 | 5081.9822 | 1767162.5
418 | 756.1639 | 1508860 | -7062.6031 | 7818.767 | 1515922.6031
418.1 | -619.665 | 1142360 | 9410.1233 | -10029.7883 | 1132949.8768
418.2 | 6220.7375 | 985850 | 4596.817 | 1623.9206 | 981253.183
418.3 | 19183.6 | 767255 | -6359.5267 | 25543.1267 | 773614.5267
418.4 | 6886.9643 | 646121 | -11162.5 | 18049.4643 | 657283.5
418.5 | 2121.9255 | 658673 | -5021.389 | 7143.3144 | 663694.389
418.6 | 3492.0476 | 546572 | 5286.6598 | -1794.6122 | 541285.3402
418.7 | -10915.1 | 580705 | 3899.7582 | -14814.8582 | 576805.2418
418.8 | -5440.8995 | 622606 | -11171.1 | 5730.2005 | 633777.1
418.9 | 751.3841 | 665084 | -205.8339 | 957.218 | 665289.834
419 | 14405.2 | 664394 | -2262.0606 | 16667.2606 | 666656.0606
419.1 | 6208.1405 | 748989 | 5977.8499 | 230.2906 | 743011.1501
419.2 | 2111.2287 | 776771 | 1851.9857 | 259.243 | 774919.0143
419.3 | -5395.4813 | 867003 | 5946.1793 | -11341.6606 | 861056.8207
419.4 | -1984.2794 | 964679 | -5015.1539 | 3030.8746 | 969694.1539
419.5 | 4182.6646 | 972257 | -2949.805 | 7132.4696 | 975206.805
419.6 | 8990.7448 | 1009820 | 9370.8135 | -380.0686 | 1000449.1865
419.7 | 15846.5 | 1115580 | 16250.5 | -404 | 1099329.5
419.8 | 11765.6 | 1193680 | 479.2714 | 11286.3286 | 1193200.7286
419.9 | 13112.8 | 1291540 | -13910.8 | 27023.6 | 1305450.8
420 | 2127.1139 | 1313310 | -12574.4 | 14701.514 | 1325884.4
420.1 | 5665.4535 | 1452100 | 6662.4695 | -997.016 | 1445437.5305
420.2 | 11083.4 | 1468420 | -8430.4172 | 19513.8172 | 1476850.4172
420.3 | 17185.4 | 1454380 | 8699.4391 | 8485.9609 | 1445680.5609
420.4 | 2119.4315 | 1575170 | -5019.7493 | 7139.1808 | 1580189.7493
420.5 | 3486.8283 | 1513380 | 10809.6 | -7322.7717 | 1502570.4
420.6 | 21292.4 | 1659610 | -14757.4 | 36049.8 | 1674367.4
420.7 | 22063.2 | 1600400 | 1177.5083 | 20885.6917 | 1599222.4917
420.8 | 15196 | 1674840 | -3619.1022 | 18815.1022 | 1678459.1022
420.9 | -4049.0883 | 1636570 | 1163.3014 | -5212.3898 | 1635406.6986
421 | 69.6324 | 1622020 | 5311.465 | -5241.8326 | 1616708.535
421.1 | 7690.5202 | 1632970 | -4333.6667 | 12024.1869 | 1637303.6667
421.2 | 5546.9054 | 1591020 | 10120.6 | -4573.6946 | 1580899.4
421.3 | 6922.0229 | 1591810 | 4607.9402 | 2314.0828 | 1587202.0598
421.4 | 21362.3 | 1603230 | -11233.4 | 32595.7 | 1614463.4
421.5 | 1442.7647 | 1516070 | 8068.1009 | -6625.3362 | 1508001.8991
421.6 | 2827.1038 | 1498350 | -12663 | 15490.1038 | 1511013
421.7 | 15910.4 | 1480270 | -7121.123 | 23031.523 | 1487391.123
421.8 | 6262.8023 | 1468850 | -17484.8 | 23747.6023 | 1486334.8
421.9 | 20767.6 | 1503810 | 4642.9866 | 16124.6134 | 1499167.0134
422 | 13170.3 | 1416010 | -2987.3771 | 16157.6771 | 1418997.3771
422.1 | 13156.4 | 1432580 | 3278.8165 | 9877.5835 | 1429301.1835
422.2 | 4899.6861 | 1252500 | -20292.3 | 25191.9861 | 1272792.3
422.3 | 6292.3783 | 1307220 | 12168.9 | -5876.5217 | 1295051.1
422.4 | -4103.3257 | 1217180 | -2281.8554 | -1821.4703 | 1219461.8554
422.5 | 20176.6 | 1192710 | 9435.6907 | 10740.9093 | 1183274.3093
422.6 | 6995.9224 | 1218990 | -11956.2 | 18952.1224 | 1230946.2
422.7 | 3526.2887 | 1165300 | -16094.3 | 19620.5887 | 1181394.3
422.8 | -1316.0673 | 1193340 | 16388 | -17704.0673 | 1176952
422.9 | 13267.7 | 1146350 | 19165.3 | -5897.6 | 1127184.7
423 | 69.1936 | 1137470 | 7429.255 | -7360.0614 | 1130040.745
423.1 | 2849.5589 | 1127660 | -5754.1032 | 8603.6621 | 1133414.1032
423.2 | -5495.3526 | 1169080 | 6735.8262 | -12231.1788 | 1162344.1738
423.3 | -3398.8045 | 1128670 | -9238.1699 | 5839.3654 | 1137908.1699
423.4 | -2699.0991 | 1217100 | 3956.5934 | -6655.6925 | 1213143.4067
423.5 | 15296.9 | 1249580 | -26656.1 | 41953 | 1276236.1
423.6 | 7688.6668 | 1176140 | -1601.9269 | 9290.5937 | 1177741.9269
423.7 | -6892.0405 | 1189650 | -7879.741 | 987.7004 | 1197529.741
423.8 | -3425.7476 | 1173060 | -5089.7676 | 1664.02 | 1178149.7676
423.9 | 8466.2397 | 1149690 | 8821.4061 | -355.1663 | 1140868.5939
424 | 5652.4088 | 1247180 | -11324.2 | 16976.6088 | 1258504.2
424.1 | 11854.9 | 1238060 | -12035.6 | 23890.5 | 1250095.6
424.2 | 7668.0031 | 1171590 | -9248.5493 | 16916.5524 | 1180838.5493
424.3 | -6226.7457 | 1246930 | -9244.8775 | 3018.1319 | 1256174.8775
424.4 | 9780.9285 | 1312950 | -12729.6 | 22510.5285 | 1325679.6
424.5 | 10633.7 | 1270000 | -2296.8733 | 12930.5733 | 1272296.8733
424.6 | -8988.7776 | 1307740 | 9537.3867 | -18526.1643 | 1298202.6133
424.7 | 16803 | 1325250 | -905.7248 | 17708.7248 | 1326155.7248
424.8 | 13290.5 | 1373880 | 3283.2204 | 10007.2796 | 1370596.7796
424.9 | 11935.8 | 1376740 | 3990.7581 | 7945.0419 | 1372749.2419
425 | -2028.0907 | 1306880 | -1611.4773 | -416.6134 | 1308491.4773
Measurement 1
(both filters in)
Integrating sphere
"scatter" NDF
"new" NDF
excitation
scattered light
Measurement 2
(remove scatter filter)
Integrating sphere
"new" NDF
excitation
scattered light
Measurement 3
(block all light using well-fitted, thick black card or similar)
Integrating sphere
excitation
BLOCK 100% of light!
detector
detector
detector

## Split Scan - PLQY
PL Quantum Yield Calculations - Split Scan method (for lower PL QY samples)
Note - Delete old data before pasting new data in. Input cells/columns are shaded green, editing other cells is not advised.
Use at your own risk & always check formulas before publishing!
Peak ranges for scatter & emission peaks (shown as red lines on graphs) | [6 empty] | OUTPUTS
|  | Scattering peak |  | Emission peak |  |  | Calculation (using as-supplied data) |  | Calculation (Jacobian Corrected)*
| Lower bound input: | 375 | nm | 540 | nm | 28120800 | Integrals and output (don't change!!)
| Upper bound input: | 385 | nm | 690 | nm | 28120800 |  | 1 |  | 1
[7 empty] | Sc_blank_in | 40054140174.9675 | Sc_blank_in | 1.5209e+13
| Neutral density filter (NDF) corrections - important! |  |  |  |  | 2217150 | Sc_sample_in | 13242157562.1361 | Sc_sample_in | 5.0276e+12
| Use the "Split Scan - NDF" sheet for the 2NDF measurement inputs (more accurate) |  |  |  |  | 2217150 | Sc_sample_out | 0 | Sc_sample_out | 0
|  | Choice (pick from list): | 2 NDF meas |  |  |  | Em_blank_in | 121428542.2167 | Em_blank_in | 75514705749.6976
[7 empty] | Em_sample_in | 248139052 | Em_sample_in | 153621422184
| Two NDF measurement (input data in Split Scan - NDF sheet) | [5 empty] | Em_sample_out | 0 | Em_sample_out | 0
|  | Filter in | 13536438.7646 |  |  |  | Sample %Absorption | 66.94% | Sample %Absorption | 66.94%
|  | Filter out | 2262493455.9301 |  |  |  | Sample Absorbance | 0.48069 | Sample Absorbance | 0.48075
[7 empty] | simple PL QY [1] | 0.0047259 | simple PL QY [1] | 0.0076713
| Input %T manually (IF ALREADY KNOWN FOR THIS EXCITATION WAVELENGTH!) | [5 empty] | Corrected PL QY [2] | -0.0030316 | Corrected PL QY [2] | -0.0049651
|  | Input NDF %T | 0.5 | % |  |  | Porres et al PL QY [3] | 0.0077575 | Porres et al PL QY [3] | 0.012636
|  | Output NDF %T (at excitation WL) = | 0.5983%
REQUIRED INPUTS - SPECTRAL DATA | [7 empty] | OPTIONAL INPUTS - OUT OF BEAM
Blank Scatter (In beam) |  | Blank Emission (In beam) |  | Sample scatter (in beam) |  | Sample emission (in beam) |  | Sample scatter (OUT of Beam) |  | Sample emission (OUT of Beam) |  | Data output quality checks!
Wavelength | Data | Wavelength | Data | Wavelength | Data | Wavelength | Data | Wavelength | Data | Wavelength | Data | Scatter graph:
370 | 54944.3 | 500 | 1116157.1429 | 370 | 12230.5 | 500 | 1213040 |  |  |  |  | The "sample in beam" scatter peak should be much lower than the "blank" scatter peak.
370.1 | 82126.4 | 501 | 1107345.2594 | 370.1 | 29492.1 | 501 | 1157920 |  |  |  |  | If these peaks are too close, the error will be large.
370.2 | 73405.5 | 502 | 1084132.6165 | 370.2 | 25133.9 | 502 | 1142020 |  |  |  |  | This means you need to either (i) increase absorbance of sample or (ii) excite at a wavelength where sample absorbs more light.
370.3 | 67672.9 | 503 | 1063935.368 | 370.3 | 30264.7 | 503 | 1181950 |  |  |  |  | If you measured "out of beam" as well, check that the "out of beam" scatter peak is significantly below the "blank" scatter peak & significantly above the "sample" peak. If not, the absorbance is too low, or there was an error with the measurement and you should NOT trust the corrected PLQY.
370.4 | 66801 | 504 | 1067086.3309 | 370.4 | 42429.2 | 504 | 1132900
370.5 | 68431.7 | 505 | 1072201.8018 | 370.5 | 27297.9 | 505 | 1131320 |  |  |  |  | Emission graph:
370.6 | 58878 | 506 | 1054108.3032 | 370.6 | 38162.2 | 506 | 1247320 |  |  |  |  | The sample in beam emission should be much higher than the "sample out of beam" emission (if measured) and the "blank emission (baseline).
370.7 | 65904.7 | 507 | 1043518.9873 | 370.7 | 36647.2 | 507 | 1151640 |  |  |  |  | The baseline (blank emission) should match well with the baseline of the sample emission.
370.8 | 56466.8 | 508 | 1031402.1739 | 370.8 | 18003.4 | 508 | 1091270 |  |  |  |  | If the sample emission is too close to the blank emission line, the PLQY is probably very low or the absorbance should be higher (see scatter plot notes)
370.9 | 73842.2 | 509 | 1016442.8312 | 370.9 | 32936.9 | 509 | 1107550
371 | 59524.7 | 510 | 1017549.0909 | 371 | 27217.5 | 510 | 1120030 |  |  |  |  | *Jacobian Correction:
371.1 | 88108.8 | 511 | 1020204.0073 | 371.1 | 9290.9761 | 511 | 1076810 |  |  |  |  | Some spectrometer data is already output as spectral photon intensity and so integrations of areas in different regions of spectra can be directly compared to compute relative photon counts. Other spectrometers output data which requires further processing using a "Jacobian transformation" (converting wavelength units to energy units). See https://pubs.acs.org/doi/10.1021/jz401508t
371.2 | 81761.2 | 512 | 1032649.635 | 371.2 | 31397.4 | 512 | 1054460 |  |  |  |  | If in doubt - measure a standard (e.g. Rhodamine 101) and confirm which output is most correct, or contact spectrometer supplier
371.3 | 79271.4 | 513 | 1004731.2614 | 371.3 | 52034.9 | 513 | 1043830
371.4 | 67401.3 | 514 | 988106.2271 | 371.4 | 24975.7 | 514 | 989895
371.5 | 106658 | 515 | 1008157.7982 | 371.5 | 43481.4 | 515 | 1081370
371.6 | 92758.5 | 516 | 1002709.5588 | 371.6 | 59059.6 | 516 | 1144800
371.7 | 74946.3 | 517 | 1040419.8895 | 371.7 | 35613.3 | 517 | 1092570
371.8 | 98143.3 | 518 | 969143.9114 | 371.8 | 29128.2 | 518 | 1021690
371.9 | 85431.1 | 519 | 948439.9261 | 371.9 | 34972.3 | 519 | 1003670
372 | 81826.6 | 520 | 954381.4815 | 372 | 39033.7 | 520 | 1044560
372.1 | 91737.8 | 521 | 958601.1132 | 372.1 | 34788.8 | 521 | 1086410
372.2 | 103916 | 522 | 929059.4796 | 372.2 | 33434.9 | 522 | 997745
372.3 | 96064.3 | 523 | 921370.5773 | 372.3 | 41179.7 | 523 | 1024550
372.4 | 100682 | 524 | 899962.6866 | 372.4 | 32615 | 524 | 1002950
372.5 | 93407.4 | 525 | 905865.4206 | 372.5 | 28335.1 | 525 | 992823
372.6 | 95237.5 | 526 | 960928.839 | 372.6 | 35407.2 | 526 | 1044820
372.7 | 92014.3 | 527 | 922592.8705 | 372.7 | 28313.3 | 527 | 925711
372.8 | 104991 | 528 | 890845.8647 | 372.8 | 44468.1 | 528 | 1040590
372.9 | 103857 | 529 | 933529.1902 | 372.9 | 37459.8 | 529 | 1033580
373 | 99084.9 | 530 | 935007.5472 | 373 | 34541 | 530 | 974426
373.1 | 116886 | 531 | 847595.4631 | 373.1 | 36696.4 | 531 | 1007510
373.2 | 118931 | 532 | 891412.8788 | 373.2 | 37425 | 532 | 927759
373.3 | 122078 | 533 | 904732.4478 | 373.3 | 40794.6 | 533 | 1002280
373.4 | 109408 | 534 | 892817.4905 | 373.4 | 65014.7 | 534 | 967064
373.5 | 117830 | 535 | 882300.9524 | 373.5 | 38759.5 | 535 | 950061
373.6 | 110111 | 536 | 869904.5802 | 373.6 | 47173.2 | 536 | 928162
373.7 | 142458 | 537 | 875586.9981 | 373.7 | 41624.4 | 537 | 942228
373.8 | 153215 | 538 | 818210.728 | 373.8 | 63395 | 538 | 1065320
373.9 | 137910 | 539 | 839497.1209 | 373.9 | 60548.8 | 539 | 918640
374 | 113043 | 540 | 813307.6923 | 374 | 53472.6 | 540 | 983776
374.1 | 177762 | 541 | 826373.7958 | 374.1 | 55521.9 | 541 | 982007
374.2 | 134696 | 542 | 813189.1892 | 374.2 | 55536.1 | 542 | 965673
374.3 | 159384 | 543 | 798967.118 | 374.3 | 53219.4 | 543 | 984145
374.4 | 153199 | 544 | 832027.1318 | 374.4 | 56592 | 544 | 909225
374.5 | 159546 | 545 | 844547.5728 | 374.5 | 63549.3 | 545 | 961219
374.6 | 185642 | 546 | 860225.6809 | 374.6 | 66357.4 | 546 | 935361
374.7 | 187623 | 547 | 808861.5984 | 374.7 | 64807.7 | 547 | 988058
374.8 | 162280 | 548 | 806062.5 | 374.8 | 66993.8 | 548 | 985789
374.9 | 183147 | 549 | 763698.6301 | 374.9 | 68469.5 | 549 | 1012790
375 | 202123 | 550 | 799015.6863 | 375 | 55805.5 | 550 | 994395
375.1 | 209014 | 551 | 806522.5933 | 375.1 | 72983.7 | 551 | 1030300
375.2 | 208875 | 552 | 811716.5354 | 375.2 | 75971.5 | 552 | 982060
375.3 | 231697 | 553 | 742311.6371 | 375.3 | 86375.9 | 553 | 995877
375.4 | 224396 | 554 | 758225.2964 | 375.4 | 70967.8 | 554 | 1032100
375.5 | 263734 | 555 | 800075.2475 | 375.5 | 103718 | 555 | 1051180
375.6 | 232570 | 556 | 795341.2698 | 375.6 | 93555.4 | 556 | 1079490
375.7 | 282510 | 557 | 754687.8728 | 375.7 | 102989 | 557 | 1052660
375.8 | 294211 | 558 | 783310.757 | 375.8 | 98833.6 | 558 | 997617
375.9 | 334739 | 559 | 771009.98 | 375.9 | 121945 | 559 | 1041580
376 | 352690 | 560 | 723400 | 376 | 163031 | 560 | 1049020
376.1 | 415592 | 561 | 739378.7575 | 376.1 | 165301 | 561 | 1106050
376.2 | 528779 | 562 | 729654.6185 | 376.2 | 267448 | 562 | 1163950
376.3 | 632525 | 563 | 728241.4487 | 376.3 | 260853 | 563 | 1108160
376.4 | 864001 | 564 | 708629.0323 | 376.4 | 383363 | 564 | 1156940
376.5 | 1055010 | 565 | 744707.0707 | 376.5 | 492163 | 565 | 1190340
376.6 | 1298140 | 566 | 720461.5385 | 376.6 | 517223 | 566 | 1255160
376.7 | 1734170 | 567 | 715237.3225 | 376.7 | 738069 | 567 | 1271900
376.8 | 2133940 | 568 | 716317.0732 | 376.8 | 919956 | 568 | 1230610
376.9 | 3037360 | 569 | 717160.8961 | 376.9 | 1326770 | 569 | 1271110
377 | 3524790 | 570 | 684889.7959 | 377 | 1552940 | 570 | 1292930
377.1 | 4866750 | 571 | 711443.7628 | 377.1 | 2074020 | 571 | 1259870
377.2 | 5717850 | 572 | 711098.3607 | 377.2 | 2497430 | 572 | 1368500
377.3 | 6557920 | 573 | 696644.7639 | 377.3 | 2734790 | 573 | 1444880
377.4 | 8988420 | 574 | 702617.284 | 377.4 | 3728290 | 574 | 1438930
377.5 | 10507200 | 575 | 697632.9897 | 377.5 | 4224830 | 575 | 1414600
377.6 | 13865200 | 576 | 691384.2975 | 377.6 | 5446180 | 576 | 1403690
377.7 | 15448500 | 577 | 654020.7039 | 377.7 | 6057640 | 577 | 1519900
377.8 | 19499500 | 578 | 674394.1909 | 377.8 | 7542230 | 578 | 1552520
377.9 | 21870200 | 579 | 725983.368 | 377.9 | 8622230 | 579 | 1538640
378 | 27428500 | 580 | 656275 | 378 | 10301100 | 580 | 1524160
378.1 | 30053700 | 581 | 679736.952 | 378.1 | 11204800 | 581 | 1586880
378.2 | 32948300 | 582 | 642221.7573 | 378.2 | 12279900 | 582 | 1663040
378.3 | 39136700 | 583 | 658067.086 | 378.3 | 14358600 | 583 | 1689640
378.4 | 42145000 | 584 | 652760.5042 | 378.4 | 15215100 | 584 | 1653100
378.5 | 49173700 | 585 | 686762.1053 | 378.5 | 17349300 | 585 | 1698260
378.6 | 52705100 | 586 | 624109.7046 | 378.6 | 18205600 | 586 | 1717690
378.7 | 59304800 | 587 | 697095.1374 | 378.7 | 20625600 | 587 | 1795140
378.8 | 62281500 | 588 | 653351.6949 | 378.8 | 21516000 | 588 | 1836210
378.9 | 65787800 | 589 | 629193.2059 | 378.9 | 22613900 | 589 | 1773250
379 | 71747100 | 590 | 674676.5957 | 379 | 24399400 | 590 | 1836920
379.1 | 73835200 | 591 | 637829.4243 | 379.1 | 24707300 | 591 | 1817150
379.2 | 79324100 | 592 | 663478.6325 | 379.2 | 26465800 | 592 | 1843140
379.3 | 81653600 | 593 | 624899.3576 | 379.3 | 26897300 | 593 | 1902400
379.4 | 84102300 | 594 | 651965.6652 | 379.4 | 27584700 | 594 | 1910690
379.5 | 85757200 | 595 | 646967.7419 | 379.5 | 27986200 | 595 | 1904510
379.6 | 87312800 | 596 | 627896.5517 | 379.6 | 28120800 | 596 | 2044150
379.7 | 87589400 | 597 | 683377.9698 | 379.7 | 27909900 | 597 | 1987290
379.8 | 87254100 | 598 | 627554.1126 | 379.8 | 27519000 | 598 | 1982400
379.9 | 86048100 | 599 | 617852.4946 | 379.9 | 27156100 | 599 | 2085790
380 | 84441200 | 600 | 628021.7391 | 380 | 26675000 | 600 | 2077850
380.1 | 81527400 | 601 | 624427.0153 | 380.1 | 25672200 | 601 | 2042260
380.2 | 78793100 | 602 | 671458.5153 | 380.2 | 24601300 | 602 | 2011920
380.3 | 74810300 | 603 | 647562.3632 | 380.3 | 23108700 | 603 | 2217150
380.4 | 72298300 | 604 | 629811.4035 | 380.4 | 22159400 | 604 | 2062240
380.5 | 68641500 | 605 | 599841.7582 | 380.5 | 21176500 | 605 | 2060980
380.6 | 63057300 | 606 | 642431.7181 | 380.6 | 19313500 | 606 | 2072550
380.7 | 59393500 | 607 | 645660.0442 | 380.7 | 18094700 | 607 | 2066090
380.8 | 52095600 | 608 | 625230.0885 | 380.8 | 15911900 | 608 | 2145560
380.9 | 49238500 | 609 | 602629.7118 | 380.9 | 14836100 | 609 | 2133380
381 | 41879400 | 610 | 635502.2222 | 381 | 12901800 | 610 | 2185390
381.1 | 38475800 | 611 | 638013.363 | 381.1 | 11797600 | 611 | 2073420
381.2 | 32867700 | 612 | 630066.9643 | 381.2 | 10084200 | 612 | 2117880
381.3 | 29650300 | 613 | 655829.9776 | 381.3 | 9062930 | 613 | 2212640
381.4 | 26492100 | 614 | 618726.4574 | 381.4 | 8095800 | 614 | 2193250
381.5 | 21293900 | 615 | 619689.8876 | 381.5 | 6733510 | 615 | 2197310
381.6 | 19159200 | 616 | 623036.036 | 381.6 | 6081130 | 616 | 2146820
381.7 | 14987800 | 617 | 649024.8307 | 381.7 | 4693800 | 617 | 2176290
381.8 | 13042400 | 618 | 625027.1493 | 381.8 | 4021120 | 618 | 2141180
381.9 | 9979210 | 619 | 660929.7052 | 381.9 | 3062500 | 619 | 2201590
382 | 8610500 | 620 | 636177.2727 | 382 | 2721550 | 620 | 2042440
382.1 | 7479450 | 621 | 650246.0137 | 382.1 | 2389690 | 621 | 2122510
382.2 | 5309410 | 622 | 769954.3379 | 382.2 | 1770870 | 622 | 2145750
382.3 | 4487210 | 623 | 784668.1922 | 382.3 | 1527360 | 623 | 2168240
382.4 | 3225740 | 624 | 763311.9266 | 382.4 | 1121350 | 624 | 2205080
382.5 | 2550660 | 625 | 686257.4713 | 382.5 | 921387 | 625 | 2057970
382.6 | 1851390 | 626 | 625580.6452 | 382.6 | 677516 | 626 | 2089240
382.7 | 1513750 | 627 | 655565.8199 | 382.7 | 624873 | 627 | 1888450
382.8 | 1131480 | 628 | 649046.2963 | 382.8 | 448824 | 628 | 1936360
382.9 | 960654 | 629 | 667874.71 | 382.9 | 381958 | 629 | 2035450
383 | 861449 | 630 | 646581.3953 | 383 | 390354 | 630 | 2068080
383.1 | 677877 | 631 | 625613.0536 | 383.1 | 326571 | 631 | 1953380
383.2 | 663431 | 632 | 604135.514 | 383.2 | 309566 | 632 | 2037820
383.3 | 611464 | 633 | 640669.7892 | 383.3 | 268032 | 633 | 1992460
383.4 | 583855 | 634 | 668563.3803 | 383.4 | 298850 | 634 | 1803910
383.5 | 561281 | 635 | 707251.7647 | 383.5 | 278793 | 635 | 1941010
383.6 | 600469 | 636 | 642537.7358 | 383.6 | 298584 | 636 | 1911300
383.7 | 633616 | 637 | 722406.6194 | 383.7 | 305671 | 637 | 1837940
383.8 | 610116 | 638 | 727483.4123 | 383.8 | 354929 | 638 | 1781150
383.9 | 636789 | 639 | 638931.1164 | 383.9 | 317812 | 639 | 1907730
384 | 740060 | 640 | 687642.8571 | 384 | 389640 | 640 | 1895090
384.1 | 725366 | 641 | 662362.7685 | 384.1 | 411123 | 641 | 1978420
384.2 | 784742 | 642 | 694440.1914 | 384.2 | 464660 | 642 | 1793210
384.3 | 789398 | 643 | 679932.8537 | 384.3 | 455189 | 643 | 1773080
384.4 | 894185 | 644 | 687153.8462 | 384.4 | 518493 | 644 | 1809290
384.5 | 977304 | 645 | 675754.2169 | 384.5 | 519072 | 645 | 1849110
384.6 | 955906 | 646 | 702357.4879 | 384.6 | 519069 | 646 | 1745220
384.7 | 975807 | 647 | 727845.0363 | 384.7 | 550438 | 647 | 1730540
384.8 | 1088980 | 648 | 773946.6019 | 384.8 | 635068 | 648 | 1545060
384.9 | 1161740 | 649 | 739309.0024 | 384.9 | 635958 | 649 | 1753780
385 | 1174180 | 650 | 736424.3902 | 385 | 618645 | 650 | 1766550
385.1 | 1256860 | 651 | 750171.1491 | 385.1 | 739269 | 651 | 1715200
385.2 | 1211460 | 652 | 717465.6863 | 385.2 | 665782 | 652 | 1668840
385.3 | 1308110 | 653 | 800422.6044 | 385.3 | 738950 | 653 | 1570520
385.4 | 1349280 | 654 | 835600.9852 | 385.4 | 740975 | 654 | 1616320
385.5 | 1355430 | 655 | 795279.0123 | 385.5 | 762736 | 655 | 1737780
385.6 | 1429310 | 656 | 800361.3861 | 385.6 | 760345 | 656 | 1654880
385.7 | 1368930 | 657 | 833950.3722 | 385.7 | 758544 | 657 | 1540390
385.8 | 1388030 | 658 | 807169.1542 | 385.8 | 762469 | 658 | 1467920
385.9 | 1369250 | 659 | 860698.2544 | 385.9 | 772014 | 659 | 1524010
386 | 1370290 | 660 | 858270 | 386 | 736163 | 660 | 1526360
386.1 | 1382140 | 661 | 893659.1479 | 386.1 | 693183 | 661 | 1521080
386.2 | 1358980 | 662 | 903402.0101 | 386.2 | 721874 | 662 | 1599170
386.3 | 1471250 | 663 | 909375.3149 | 386.3 | 705193 | 663 | 1479550
386.4 | 1406450 | 664 | 916944.4444 | 386.4 | 690600 | 664 | 1676580
386.5 | 1429340 | 665 | 924744.3038 | 386.5 | 668620 | 665 | 1532820
386.6 | 1337310 | 666 | 1030451.7766 | 386.6 | 692726 | 666 | 1498610
386.7 | 1363490 | 667 | 1027684.4784 | 386.7 | 649856 | 667 | 1655170
386.8 | 1315290 | 668 | 934923.4694 | 386.8 | 664318 | 668 | 1586340
386.9 | 1270280 | 669 | 1051319.6931 | 386.9 | 664745 | 669 | 1592950
387 | 1228840 | 670 | 1010533.3333 | 387 | 608629 | 670 | 1670730
387.1 | 1276920 | 671 | 1132102.8278 | 387.1 | 599291 | 671 | 1526020
387.2 | 1264620 | 672 | 1147247.4227 | 387.2 | 624770 | 672 | 1623360
387.3 | 1276940 | 673 | 1210186.0465 | 387.3 | 638443 | 673 | 1636650
387.4 | 1220570 | 674 | 1152352.3316 | 387.4 | 571064 | 674 | 1476980
387.5 | 1275320 | 675 | 1184218.1818 | 387.5 | 618104 | 675 | 1579580
387.6 | 1172970 | 676 | 1240015.625 | 387.6 | 593088 | 676 | 1581910
387.7 | 1201570 | 677 | 1225618.799 | 387.7 | 575323 | 677 | 1595510
387.8 | 1201340 | 678 | 1207560.2094 | 387.8 | 603285 | 678 | 1450250
387.9 | 1224170 | 679 | 1301716.5354 | 387.9 | 574814 | 679 | 1658260
388 | 1239980 | 680 | 1287547.3684 | 388 | 601425 | 680 | 1675490
388.1 | 1226050 | 681 | 1268749.3404 | 388.1 | 580611 | 681 | 1678720
388.2 | 1199180 | 682 | 1286148.1481 | 388.2 | 565763 | 682 | 1564070
388.3 | 1254670 | 683 | 1454180.3714 | 388.3 | 553953 | 683 | 1633480
388.4 | 1243260 | 684 | 1431079.7872 | 388.4 | 556526 | 684 | 1685880
388.5 | 1231780 | 685 | 1474618.6667 | 388.5 | 584254 | 685 | 1558650
388.6 | 1192780 | 686 | 1509454.5455 | 388.6 | 570920 | 686 | 1805250
388.7 | 1223050 | 687 | 1565930.2949 | 388.7 | 604545 | 687 | 1589800
388.8 | 1244210 | 688 | 1662451.6129 | 388.8 | 556238 | 688 | 1674770
388.9 | 1146490 | 689 | 1641649.5957 | 388.9 | 580809 | 689 | 1652070
389 | 1164210 | 690 | 1583491.8919 | 389 | 554594 | 690 | 1683260
389.1 | 1185210 | 691 | 1784905.1491 | 389.1 | 569460 | 691 | 1721140
389.2 | 1142270 | 692 | 1761532.6087 | 389.2 | 579505 | 692 | 1639540
389.3 | 1201130 | 693 | 1751035.4223 | 389.3 | 559556 | 693 | 1646580
389.4 | 1198010 | 694 | 1841251.3661 | 389.4 | 575137 | 694 | 1816760
389.5 | 1163290 | 695 | 1894536.9863 | 389.5 | 585711 | 695 | 1754640
389.6 | 1112110 | 696 | 1838719.7802 | 389.6 | 584566 | 696 | 1714380
389.7 | 1189440 | 697 | 1880573.0028 | 389.7 | 548555 | 697 | 1747180
389.8 | 1141910 | 698 | 2076972.3757 | 389.8 | 557689 | 698 | 1816670
389.9 | 1159380 | 699 | 1923800.554 | 389.9 | 587702 | 699 | 1795880
390 | 1119130 | 700 | 2045344.4444 | 390 | 602092 | 700 | 1764870

## Single Scan
PL Quantum Yield Calculations - Single Scan method (for high PL QY samples)
Note - Delete old data before pasting new data in. Input cells/columns are shaded green, editing other cells is not advised.
Use at your own risk & always check formulas before publishing!
INPUTS | [5 empty] | OUTPUTS |  |  |  | Data output quality checks!
|  | Scattering peak |  | Emission peak | [5 empty] | Scatter graph:
| Lower bound input: | 470 | nm | 525 | nm | Calculation (using as-supplied data) |  | Graphs (linear & log-lin plots) |  | The "sample in beam" scatter peak should be much lower than the "blank" scatter peak.
| Upper bound input: | 490 | nm | 670 | nm | Parameter | Calculation |  |  | If these peaks are too close, the error will be large.
[6 empty] | Sc_blank_in | 69100161.5506 |  | 1702000 | This means you need to either (i) increase absorbance of sample or (ii) excite at a wavelength where sample absorbs more light.
Blank in beam |  | Sample in beam |  | Sample out of Beam (optional!) |  | Sc_sample_in | 52448109.6641 |  | 1702000 | If you measured "out of beam" as well, check that the "out of beam" scatter peak is significantly below the "blank" scatter peak & significantly above the "sample" peak. If not, the absorbance is too low, or there was an error with the measurement and you should NOT trust the corrected PLQY.
Wavelength | Data | Wavelength | Data | Wavelength | Data | Sc_sample_out | 0
460 | 1698.1628 | 460 | 1378.318 |  |  | Em_blank_in | 213121.2681 |  |  | Emission graph:
460.2 | 1795.6924 | 460.2 | 1347.2427 |  |  | Em_sample_in | 10843769.0952 |  | 1304630 | The sample in beam emission should be much higher than the "sample out of beam" emission (if measured) and the "blank emission (baseline).
460.4 | 1844.4157 | 460.4 | 1299.2775 |  |  | Em_sample_out | 0 |  | 1304630 | The baseline (blank emission) should match well with the baseline of the sample emission.
460.6 | 1745.2733 | 460.6 | 1394.6798 |  |  | Sample %Absorption | 0.24098 |  |  | If the sample emission is too close to the blank emission line, the PLQY is probably very low or the absorbance should be higher (see scatter plot notes)
460.8 | 1991.699 | 460.8 | 1430.0579 |  |  | Sample Absorbance | 0.11975
461 | 2017.9159 | 461 | 1419.6342 |  |  | simple PL QY [1] | 63.84% |  |  | *Jacobian Correction:
461.2 | 2003.4847 | 461.2 | 1537.0552 |  |  | Corrected PL QY [2] | -0.3084% |  |  | Some spectrometer data is already output as spectral photon intensity and so integrations of areas in different regions of spectra can be directly compared to compute relative photon counts. Other spectrometers output data which requires further processing using a "Jacobian transformation" (converting wavelength units to energy units). See https://pubs.acs.org/doi/10.1021/jz401508t
461.4 | 2077.5187 | 461.4 | 1606.3504 |  |  | Porres et al PL QY [3] | 64.15% |  |  | If in doubt - measure a standard (e.g. Rhodamine 101) and confirm which output is most correct, or contact spectrometer supplier
461.6 | 2026.9814 | 461.6 | 1579.8865
461.8 | 2088.765 | 461.8 | 1683.8596 |  |  | Calculation (Jacobian corrected*)
462 | 2335.3829 | 462 | 1660.7316 |  |  | Parameter | Calculation
462.2 | 2259.0334 | 462.2 | 1728.5819 |  |  | Sc_blank_in | 33187300885.6164
462.4 | 2516.0366 | 462.4 | 1816.6033 |  |  | Sc_sample_in | 25191354774.7216
462.6 | 2573.7307 | 462.6 | 1734.7145 |  |  | Sc_sample_out | 0
462.8 | 2324.784 | 462.8 | 1800.0956 |  |  | Em_blank_in | 124843426.4443
463 | 2389.1514 | 463 | 1729.3605 |  |  | Em_sample_in | 6262734302.3352
463.2 | 2506.1223 | 463.2 | 1722.1406 |  |  | Em_sample_out | 0
463.4 | 2339.3798 | 463.4 | 1777.9512 |  |  | Sample %Absorption | 0.24093
463.6 | 2434.9324 | 463.6 | 1768.6897 |  |  | Sample Absorbance | 0.11972
463.8 | 2583.5889 | 463.8 | 1691.4698 |  |  | simple PL QY [1] | 76.76%
464 | 2341.7713 | 464 | 1809.3091 |  |  | Corrected PL QY [2] | -0.3762%
464.2 | 2401.5759 | 464.2 | 1644.4235 |  |  | Porres et al PL QY [3] | 77.14%
464.4 | 2422.4101 | 464.4 | 1698.3597
464.6 | 2422.243 | 464.6 | 1805.9578
464.8 | 2548.3266 | 464.8 | 1739.0512
465 | 2494.4124 | 465 | 1941.9506
465.2 | 2523.1883 | 465.2 | 1815.308
465.4 | 2557.8489 | 465.4 | 1941.6125
465.6 | 2759.6949 | 465.6 | 2002.4193
465.8 | 2898.7683 | 465.8 | 1900.1336
466 | 2951.5444 | 466 | 2148.6341
466.2 | 3120.4917 | 466.2 | 2328.8617
466.4 | 3208.112 | 466.4 | 2526.7043
466.6 | 3513.6017 | 466.6 | 2487.2018
466.8 | 4071.9766 | 466.8 | 2781.2892
467 | 4582.6043 | 467 | 3203.1281
467.2 | 4742.5819 | 467.2 | 3389.3114
467.4 | 4604.3891 | 467.4 | 3238.767
467.6 | 4715.0961 | 467.6 | 3668.6429
467.8 | 4473.2662 | 467.8 | 3573.3575
468 | 4696.0642 | 468 | 3369.9867
468.2 | 4216.6723 | 468.2 | 3306.0185
468.4 | 4326.6551 | 468.4 | 3329.432
468.6 | 4369.7453 | 468.6 | 3305.3026
468.8 | 4479.3003 | 468.8 | 3105.617
469 | 4799.9301 | 469 | 3284.7194
469.2 | 4675.9969 | 469.2 | 3286.4276
469.4 | 4535.098 | 469.4 | 3311.8178
469.6 | 4495.9439 | 469.6 | 3479.8601
469.8 | 5058.195 | 469.8 | 3955.7373
470 | 5722.0884 | 470 | 4180.2405
470.2 | 6594.3425 | 470.2 | 4603.3632
470.4 | 8293.7531 | 470.4 | 6141.2302
470.6 | 11387.3 | 470.6 | 8680.5026
470.8 | 16955.4 | 470.8 | 13579.5
471 | 25872.4 | 471 | 19481.8
471.2 | 36981.9 | 471.2 | 27733.2
471.4 | 50148.3 | 471.4 | 38751.8
471.6 | 66721.6 | 471.6 | 51380.9
471.8 | 87639.6 | 471.8 | 66548.5
472 | 111089 | 472 | 83573.5
472.2 | 140145 | 472.2 | 104511
472.4 | 178205 | 472.4 | 133554
472.6 | 228576 | 472.6 | 168993
472.8 | 267155 | 472.8 | 200071
473 | 337062 | 473 | 252051
473.2 | 402829 | 473.2 | 300966
473.4 | 414963 | 473.4 | 305884
473.6 | 406927 | 473.6 | 303234
473.8 | 411939 | 473.8 | 305041
474 | 423675 | 474 | 317237
474.2 | 444614 | 474.2 | 328501
474.4 | 472896 | 474.4 | 347943
474.6 | 502792 | 474.6 | 371181
474.8 | 531778 | 474.8 | 395567
475 | 568006 | 475 | 417957
475.2 | 587573 | 475.2 | 436330
475.4 | 625240 | 475.4 | 462925
475.6 | 655594 | 475.6 | 488704
475.8 | 691380 | 475.8 | 514084
476 | 722596 | 476 | 538273
476.2 | 759101 | 476.2 | 563636
476.4 | 792406 | 476.4 | 591640
476.6 | 830350 | 476.6 | 618642
476.8 | 859019 | 476.8 | 641905
477 | 900339 | 477 | 671087
477.2 | 928875 | 477.2 | 692969
477.4 | 970341 | 477.4 | 729373
477.6 | 1000030 | 477.6 | 754955
477.8 | 1025320 | 477.8 | 774811
478 | 1068530 | 478 | 805188
478.2 | 1107520 | 478.2 | 833262
478.4 | 1147490 | 478.4 | 871419
478.6 | 1187800 | 478.6 | 901360
478.8 | 1235550 | 478.8 | 936641
479 | 1277860 | 479 | 964156
479.2 | 1330690 | 479.2 | 1008040
479.4 | 1363810 | 479.4 | 1038490
479.6 | 1408830 | 479.6 | 1070490
479.8 | 1451480 | 479.8 | 1107030
480 | 1512380 | 480 | 1153510
480.2 | 1581170 | 480.2 | 1207040
480.4 | 1632100 | 480.4 | 1245280
480.6 | 1697030 | 480.6 | 1300620
480.8 | 1702000 | 480.8 | 1304630
481 | 1660980 | 481 | 1276840
481.2 | 1585490 | 481.2 | 1217220
481.4 | 1512530 | 481.4 | 1158290
481.6 | 1448640 | 481.6 | 1107270
481.8 | 1391450 | 481.8 | 1067180
482 | 1351030 | 482 | 1037300
482.2 | 1329180 | 482.2 | 1016020
482.4 | 1313840 | 482.4 | 1005980
482.6 | 1305450 | 482.6 | 1001580
482.8 | 1280950 | 482.8 | 978734
483 | 1223190 | 483 | 940444
483.2 | 1189020 | 483.2 | 908362
483.4 | 1125000 | 483.4 | 861919
483.6 | 1064530 | 483.6 | 814817
483.8 | 1010730 | 483.8 | 779986
484 | 964952 | 484 | 741649
484.2 | 922570 | 484.2 | 705781
484.4 | 871646 | 484.4 | 668191
484.6 | 821267 | 484.6 | 630439
484.8 | 770770 | 484.8 | 593243
485 | 723557 | 485 | 554061
485.2 | 671227 | 485.2 | 510519
485.4 | 625287 | 485.4 | 480789
485.6 | 578343 | 485.6 | 441146
485.8 | 533831 | 485.8 | 407618
486 | 506429 | 486 | 384803
486.2 | 463054 | 486.2 | 353796
486.4 | 420587 | 486.4 | 321996
486.6 | 383928 | 486.6 | 290292
486.8 | 341735 | 486.8 | 261505
487 | 302415 | 487 | 229598
487.2 | 265297 | 487.2 | 199263
487.4 | 225121 | 487.4 | 170866
487.6 | 187507 | 487.6 | 143577
487.8 | 150387 | 487.8 | 112209
488 | 115836 | 488 | 87394.9
488.2 | 83883.9 | 488.2 | 62188
488.4 | 57479.6 | 488.4 | 41701.4
488.6 | 35810.2 | 488.6 | 25549
488.8 | 25124.6 | 488.8 | 18300.6
489 | 15697.4 | 489 | 11328.8
489.2 | 10294.5 | 489.2 | 7447.6159
489.4 | 6929.0626 | 489.4 | 5178.7543
489.6 | 5861.2143 | 489.6 | 4542.8007
489.8 | 5070.0211 | 489.8 | 3838.5741
490 | 4883.3685 | 490 | 3520.6828
490.2 | 4620.2916 | 490.2 | 3524.9025
490.4 | 4370.4002 | 490.4 | 3046.4021
490.6 | 4203.6501 | 490.6 | 3700.1924
490.8 | 4311.5784 | 490.8 | 3225.1861
491 | 3983.6709 | 491 | 3222.3108
491.2 | 4482.9047 | 491.2 | 3467.2979
491.4 | 4325.5019 | 491.4 | 3735.3135
491.6 | 4407.9589 | 491.6 | 3283.7426
491.8 | 4286.2507 | 491.8 | 3211.0149
492 | 4076.6235 | 492 | 3196.8501
492.2 | 3975.8517 | 492.2 | 3253.7928
492.4 | 3896.6065 | 492.4 | 3317.0325
492.6 | 3924.5648 | 492.6 | 2964.7087
492.8 | 3576.1631 | 492.8 | 2888.4661
493 | 3237.9144 | 493 | 2873.2919
493.2 | 3162.9391 | 493.2 | 2373.2955
493.4 | 3259.3256 | 493.4 | 2550.9458
493.6 | 3050.1499 | 493.6 | 2357.36
493.8 | 2801.6696 | 493.8 | 2269.4636
494 | 2642.7539 | 494 | 2115.1438
494.2 | 2771.0547 | 494.2 | 2380.6402
494.4 | 2687.3474 | 494.4 | 2049.5256
494.6 | 2598.231 | 494.6 | 1966.5136
494.8 | 2625.8296 | 494.8 | 2083.0348
495 | 2499.7773 | 495 | 1973.7182
495.2 | 2447.5615 | 495.2 | 2030.2767
495.4 | 2369.1486 | 495.4 | 1965.8875
495.6 | 2439.8012 | 495.6 | 1813.5648
495.8 | 2119.083 | 495.8 | 1680.4492
496 | 2140.9122 | 496 | 1812.9942
496.2 | 2200.4907 | 496.2 | 1851.6258
496.4 | 2210.4321 | 496.4 | 1531.3586
496.6 | 2132.3117 | 496.6 | 1785.803
496.8 | 1854.0577 | 496.8 | 1652.599
497 | 1901.184 | 497 | 1735.9743
497.2 | 1845.624 | 497.2 | 1421.6663
497.4 | 1897.9577 | 497.4 | 1563.0217
497.6 | 1964.9369 | 497.6 | 1298.0613
497.8 | 1964.1597 | 497.8 | 1468.695
498 | 1812.9831 | 498 | 1363.3896
498.2 | 1914.4706 | 498.2 | 1425.3993
498.4 | 1671.7181 | 498.4 | 1489.4374
498.6 | 1761.1774 | 498.6 | 1284.7561
498.8 | 1807.9957 | 498.8 | 1438.9809
499 | 1609.2503 | 499 | 1453.8742
499.2 | 1846.5304 | 499.2 | 1053.2404
499.4 | 1656.8383 | 499.4 | 1213.7329
499.6 | 1708.6925 | 499.6 | 1242.2907
499.8 | 1386.9702 | 499.8 | 1183.6986
500 | 1377.7972 | 500 | 1166.2392
500.2 | 1451.1963 | 500.2 | 1237.8235
500.4 | 1415.8892 | 500.4 | 1178.2822
500.6 | 1334.6087 | 500.6 | 1102.9034
500.8 | 1439.6792 | 500.8 | 1263.939
501 | 1451.377 | 501 | 1064.1915
501.2 | 1370.0936 | 501.2 | 1118.5566
501.4 | 1251.2672 | 501.4 | 1090.5621
501.6 | 1264.289 | 501.6 | 1093.2062
501.8 | 1306.883 | 501.8 | 936.8662
502 | 1399.3284 | 502 | 1030.9667
502.2 | 1427.488 | 502.2 | 1144.9396
502.4 | 1475.8551 | 502.4 | 969.5741
502.6 | 1188.1955 | 502.6 | 997.7109
502.8 | 1378.286 | 502.8 | 1213.0602
503 | 1101.4285 | 503 | 959.1894
503.2 | 1217.7424 | 503.2 | 1005.8157
503.4 | 1120.7117 | 503.4 | 1071.4886
503.6 | 1274.8303 | 503.6 | 1040.3332
503.8 | 1224.8107 | 503.8 | 1072.3782
504 | 1149.0372 | 504 | 872.6128
504.2 | 1324.5681 | 504.2 | 979.5593
504.4 | 1345.8456 | 504.4 | 907.5238
504.6 | 1152.2328 | 504.6 | 938.174
504.8 | 1044.2262 | 504.8 | 921.9971
505 | 1070.2342 | 505 | 904.4448
505.2 | 1261.7221 | 505.2 | 907.8569
505.4 | 1112.2507 | 505.4 | 907.8135
505.6 | 1123.1133 | 505.6 | 888.2845
505.8 | 1157.1626 | 505.8 | 943.4393
506 | 1222.7576 | 506 | 949.5454
506.2 | 1176.0469 | 506.2 | 814.7577
506.4 | 976.9589 | 506.4 | 881.7579
506.6 | 1034.6178 | 506.6 | 837.4288
506.8 | 942.8827 | 506.8 | 943.6338
507 | 1131.2544 | 507 | 768.2196
507.2 | 1041.7424 | 507.2 | 930.5156
507.4 | 1092.1829 | 507.4 | 801.1083
507.6 | 950.8663 | 507.6 | 713.6113
507.8 | 1071.8702 | 507.8 | 865.093
508 | 856.8342 | 508 | 637.3694
508.2 | 1018.6349 | 508.2 | 762.2301
508.4 | 997.2059 | 508.4 | 955.0819
508.6 | 1037.5825 | 508.6 | 932.295
508.8 | 939.7118 | 508.8 | 685.0018
509 | 988.009 | 509 | 809.6499
509.2 | 856.4083 | 509.2 | 696.4135
509.4 | 831.5434 | 509.4 | 801.7367
509.6 | 1113.1448 | 509.6 | 738.4793
509.8 | 890.6652 | 509.8 | 844.6129
510 | 720.1334 | 510 | 859.5953
510.2 | 896.4383 | 510.2 | 628.3531
510.4 | 942.4933 | 510.4 | 699.3402
510.6 | 848.5757 | 510.6 | 751.6409
510.8 | 910.5357 | 510.8 | 686.9869
511 | 1100.2149 | 511 | 758.8116
511.2 | 917.9141 | 511.2 | 721.4148
511.4 | 1009.5656 | 511.4 | 787.6924
511.6 | 971.7521 | 511.6 | 624.1189
511.8 | 820.1585 | 511.8 | 690.4564
512 | 828.1075 | 512 | 801.0108
512.2 | 1065.4287 | 512.2 | 611.0325
512.4 | 936.4296 | 512.4 | 592.0181
512.6 | 823.2364 | 512.6 | 802.8061
512.8 | 813.2308 | 512.8 | 782.8071
513 | 829.9789 | 513 | 831.4214
513.2 | 856.2759 | 513.2 | 627.5685
513.4 | 758.1043 | 513.4 | 683.5368
513.6 | 857.8099 | 513.6 | 670.5865
513.8 | 918.0656 | 513.8 | 659.7933
514 | 811.6672 | 514 | 694.5501
514.2 | 745.2255 | 514.2 | 770.1188
514.4 | 699.6292 | 514.4 | 678.0419
514.6 | 791.4081 | 514.6 | 704.9658
514.8 | 837.4035 | 514.8 | 583.6283
515 | 836.1249 | 515 | 569.0627
515.2 | 777.1904 | 515.2 | 587.6324
515.4 | 777.6997 | 515.4 | 717.419
515.6 | 839.3213 | 515.6 | 711.0757
515.8 | 747.0432 | 515.8 | 674.5849
516 | 738.3802 | 516 | 647.8248
516.2 | 795.6316 | 516.2 | 595.5036
516.4 | 737.9706 | 516.4 | 792.6614
516.6 | 725.1188 | 516.6 | 645.219
516.8 | 759.9676 | 516.8 | 727.8506
517 | 694.3852 | 517 | 769.202
517.2 | 765.2271 | 517.2 | 603.6585
517.4 | 793.288 | 517.4 | 541.5085
517.6 | 806.8598 | 517.6 | 657.5805
517.8 | 769.0408 | 517.8 | 685.7083
518 | 643.707 | 518 | 720.416
518.2 | 711.8726 | 518.2 | 688.0505
518.4 | 620.1229 | 518.4 | 641.423
518.6 | 803.376 | 518.6 | 492.2143
518.8 | 646.8563 | 518.8 | 676.898
519 | 648.0569 | 519 | 574.4403
519.2 | 796.7929 | 519.2 | 522.9906
519.4 | 891.3475 | 519.4 | 734.5723
519.6 | 740.3625 | 519.6 | 513.4688
519.8 | 685.6072 | 519.8 | 591.3167
520 | 761.8225 | 520 | 703.7089
520.2 | 643.3952 | 520.2 | 695.4451
520.4 | 773.7881 | 520.4 | 637.3817
520.6 | 777.9898 | 520.6 | 584.086
520.8 | 580.0154 | 520.8 | 749.8377
521 | 571.4019 | 521 | 642.3324
521.2 | 829.0495 | 521.2 | 551.5701
521.4 | 577.3148 | 521.4 | 682.0726
521.6 | 588.1675 | 521.6 | 674.4144
521.8 | 686.7776 | 521.8 | 674.6456
522 | 598.8305 | 522 | 650.2591
522.2 | 647.6244 | 522.2 | 743.4126
522.4 | 643.2032 | 522.4 | 563.226
522.6 | 758.9183 | 522.6 | 666.9305
522.8 | 716.7036 | 522.8 | 585.0601
523 | 608.3634 | 523 | 708.859
523.2 | 633.7101 | 523.2 | 701.0215
523.4 | 639.203 | 523.4 | 665.8192
523.6 | 558.9809 | 523.6 | 713.9403
523.8 | 661.3275 | 523.8 | 563.8541
524 | 652.3133 | 524 | 592.65
524.2 | 529.8872 | 524.2 | 632.9912
524.4 | 495.2791 | 524.4 | 706.7107
524.6 | 655.3528 | 524.6 | 711.141
524.8 | 452.6776 | 524.8 | 638.3814
525 | 580.6967 | 525 | 609.7915
525.2 | 653.4645 | 525.2 | 673.8402
525.4 | 674.723 | 525.4 | 749.1266
525.6 | 597.5376 | 525.6 | 645.8088
525.8 | 638.1307 | 525.8 | 656.0838
526 | 693.1198 | 526 | 697.5547
526.2 | 613.3157 | 526.2 | 596.5509
526.4 | 599.4612 | 526.4 | 726.9043
526.6 | 590.4239 | 526.6 | 721.5479
526.8 | 585.8753 | 526.8 | 813.9931
527 | 545.426 | 527 | 743.9226
527.2 | 576.0777 | 527.2 | 889.3893
527.4 | 516.9882 | 527.4 | 920.8478
527.6 | 553.1833 | 527.6 | 831.6334
527.8 | 722.1765 | 527.8 | 772.8224
528 | 552.9404 | 528 | 829.8607
528.2 | 722.8186 | 528.2 | 961.984
528.4 | 479.2253 | 528.4 | 931.8632
528.6 | 610.7654 | 528.6 | 953.724
528.8 | 581.1161 | 528.8 | 965.5471
529 | 526.6 | 529 | 976.4805
529.2 | 506.8671 | 529.2 | 1057.543
529.4 | 517.2388 | 529.4 | 950.7917
529.6 | 557.7598 | 529.6 | 876.8683
529.8 | 594.0518 | 529.8 | 1037.2414
530 | 705.3433 | 530 | 825.072
530.2 | 468.725 | 530.2 | 1034.6281
530.4 | 561.0299 | 530.4 | 1066.5876
530.6 | 601.8613 | 530.6 | 1111.7473
530.8 | 627.1624 | 530.8 | 1068.0849
531 | 587.7713 | 531 | 1207.4874
531.2 | 587.4125 | 531.2 | 1059.8456
531.4 | 725.8181 | 531.4 | 1141.52
531.6 | 661.7276 | 531.6 | 1141.1204
531.8 | 682.3529 | 531.8 | 1206.082
532 | 432.5446 | 532 | 1192.136
532.2 | 542.0271 | 532.2 | 1216.2713
532.4 | 433.5716 | 532.4 | 1299.0483
532.6 | 635.9477 | 532.6 | 1546.8435
532.8 | 664.8846 | 532.8 | 1221.3656
533 | 486.943 | 533 | 1375.1917
533.2 | 398.6278 | 533.2 | 1319.0935
533.4 | 517.6528 | 533.4 | 1525.2227
533.6 | 711.8536 | 533.6 | 1600.2765
533.8 | 467.0476 | 533.8 | 1796.9826
534 | 576.2541 | 534 | 1739.2452
534.2 | 550.4044 | 534.2 | 1506.7518
534.4 | 479.0346 | 534.4 | 1905.8862
534.6 | 510.159 | 534.6 | 1780.9021
534.8 | 444.6048 | 534.8 | 1780.7919
535 | 542.8844 | 535 | 1877.5239
535.2 | 471.4246 | 535.2 | 1998.7815
535.4 | 601.0327 | 535.4 | 2099.9294
535.6 | 382.5509 | 535.6 | 1948.5817
535.8 | 637.2909 | 535.8 | 2142.7901
536 | 483.0126 | 536 | 2337.1299
536.2 | 589.6171 | 536.2 | 2223.9621
536.4 | 458.6722 | 536.4 | 2160.1205
536.6 | 490.5348 | 536.6 | 2464.0569
536.8 | 668.1358 | 536.8 | 2296.8331
537 | 556.9513 | 537 | 2438.1163
537.2 | 569.8993 | 537.2 | 2541.1024
537.4 | 597.268 | 537.4 | 2699.0649
537.6 | 424.1822 | 537.6 | 2579.2172
537.8 | 502.8924 | 537.8 | 2733.2936
538 | 456.0648 | 538 | 3018.1379
538.2 | 483.0155 | 538.2 | 3135.8582
538.4 | 424.4772 | 538.4 | 3124.3932
538.6 | 352.3298 | 538.6 | 3005.7451
538.8 | 483.8825 | 538.8 | 3774.22
539 | 474.5162 | 539 | 3502.7302
539.2 | 473.9174 | 539.2 | 3690.5051
539.4 | 474.1454 | 539.4 | 3705.4057
539.6 | 576.0331 | 539.6 | 3772.8474
539.8 | 555.7765 | 539.8 | 4195.9616
540 | 541.2632 | 540 | 4246.0654
540.2 | 419.111 | 540.2 | 3991.0799
540.4 | 509.9661 | 540.4 | 4161.5536
540.6 | 429.5066 | 540.6 | 4052.6435
540.8 | 345.4825 | 540.8 | 4569.2634
541 | 435.716 | 541 | 5287.915
541.2 | 501.2448 | 541.2 | 5024.0858
541.4 | 384.8822 | 541.4 | 5234.5366
541.6 | 525.1319 | 541.6 | 5277.6729
541.8 | 504.1736 | 541.8 | 5410.9198
542 | 494.1209 | 542 | 5318.4738
542.2 | 475.7057 | 542.2 | 5682.0371
542.4 | 423.5729 | 542.4 | 5829.2803
542.6 | 552.415 | 542.6 | 5927.7817
542.8 | 251.9523 | 542.8 | 5900.5359
543 | 517.6312 | 543 | 6165.6737
543.2 | 491.9606 | 543.2 | 7036.1173
543.4 | 491.7212 | 543.4 | 6645.7179
543.6 | 292.4606 | 543.6 | 6821.3784
543.8 | 221.2336 | 543.8 | 7398.9981
544 | 531.9143 | 544 | 7295.7598
544.2 | 281.4309 | 544.2 | 7363.462
544.4 | 303.6386 | 544.4 | 7775.0259
544.6 | 402.1824 | 544.6 | 8033.901
544.8 | 337.7854 | 544.8 | 7911.2869
545 | 359.4857 | 545 | 8223.638
545.2 | 404.3143 | 545.2 | 8926.2038
545.4 | 317.0873 | 545.4 | 8542.1575
545.6 | 389.2348 | 545.6 | 8973.5857
545.8 | 350.6339 | 545.8 | 9138.533
546 | 472.162 | 546 | 9293.0344
546.2 | 340.4174 | 546.2 | 9811.3677
546.4 | 295.975 | 546.4 | 9916.5058
546.6 | 263.1358 | 546.6 | 10095.4
546.8 | 352.0556 | 546.8 | 10394.1
547 | 346.9573 | 547 | 11383.6
547.2 | 413.5329 | 547.2 | 10946.1
547.4 | 359.3131 | 547.4 | 11785
547.6 | 442.2939 | 547.6 | 11920.3
547.8 | 231.8522 | 547.8 | 11711.5
548 | 398.8273 | 548 | 12279.7
548.2 | 310.7168 | 548.2 | 12893.1
548.4 | 539.4327 | 548.4 | 13635.4
548.6 | 306.2892 | 548.6 | 13579
548.8 | 344.7821 | 548.8 | 13563.8
549 | 496.4673 | 549 | 14178.7
549.2 | 356.4874 | 549.2 | 14966.5
549.4 | 301.5421 | 549.4 | 15832.5
549.6 | 352.6399 | 549.6 | 15122
549.8 | 409.5748 | 549.8 | 15894.2
550 | 358.988 | 550 | 15722
550.2 | 420.9886 | 550.2 | 15832.5
550.4 | 348.4641 | 550.4 | 17043.7
550.6 | 354.4759 | 550.6 | 16601.6
550.8 | 473.4857 | 550.8 | 17910
551 | 532.2071 | 551 | 18031.8
551.2 | 390.1055 | 551.2 | 18127.4
551.4 | 361.4694 | 551.4 | 18645.2
551.6 | 379.6163 | 551.6 | 19508.3
551.8 | 272.4874 | 551.8 | 20451.9
552 | 374.0906 | 552 | 20440.9
552.2 | 237.8148 | 552.2 | 20290
552.4 | 324.156 | 552.4 | 21119.1
552.6 | 474.0509 | 552.6 | 21483.6
552.8 | 295.8229 | 552.8 | 22171.1
553 | 279.8642 | 553 | 21332.1
553.2 | 296.7767 | 553.2 | 22039.9
553.4 | 435.5202 | 553.4 | 23344.2
553.6 | 383.2334 | 553.6 | 23568
553.8 | 372.7385 | 553.8 | 24091
554 | 378.7921 | 554 | 24551.2
554.2 | 414.0453 | 554.2 | 25334.1
554.4 | 443.3066 | 554.4 | 25953.3
554.6 | 340.1361 | 554.6 | 26132.2
554.8 | 229.762 | 554.8 | 27301.2
555 | 359.0618 | 555 | 26486.2
555.2 | 277.079 | 555.2 | 28218
555.4 | 359.4203 | 555.4 | 28358.3
555.6 | 266.0957 | 555.6 | 27729.5
555.8 | 213.9915 | 555.8 | 28717.2
556 | 319.3427 | 556 | 29409
556.2 | 302.0487 | 556.2 | 30609.3
556.4 | 378.1605 | 556.4 | 31023.1
556.6 | 384.4656 | 556.6 | 31594.1
556.8 | 291.2443 | 556.8 | 31146.6
557 | 156.8386 | 557 | 32702.2
557.2 | 403.7997 | 557.2 | 32087.7
557.4 | 381.8869 | 557.4 | 33598.5
557.6 | 293.2768 | 557.6 | 33433.3
557.8 | 305.7488 | 557.8 | 34263.8
558 | 306.3729 | 558 | 35149.4
558.2 | 305.8906 | 558.2 | 34638.9
558.4 | 348.1479 | 558.4 | 35305.7
558.6 | 277.0076 | 558.6 | 36445.4
558.8 | 253.2735 | 558.8 | 36786.1
559 | 218.0682 | 559 | 37527.4
559.2 | 402.701 | 559.2 | 37107.2
559.4 | 362.1873 | 559.4 | 37616.4
559.6 | 243.0165 | 559.6 | 38856.5
559.8 | 231.0829 | 559.8 | 38317.7
560 | 333.6308 | 560 | 39620.1
560.2 | 364.4327 | 560.2 | 40033.8
560.4 | 256.6308 | 560.4 | 40370.2
560.6 | 370.7043 | 560.6 | 41183.7
560.8 | 261.9309 | 560.8 | 41058.3
561 | 408.3478 | 561 | 41891.7
561.2 | 251.7165 | 561.2 | 42297.1
561.4 | 452.3656 | 561.4 | 42408.6
561.6 | 367.6723 | 561.6 | 43069
561.8 | 415.0957 | 561.8 | 42813.1
562 | 373.6646 | 562 | 43438.6
562.2 | 331.8486 | 562.2 | 44473.3
562.4 | 167.8132 | 562.4 | 44216
562.6 | 405.6824 | 562.6 | 45286.4
562.8 | 211.2501 | 562.8 | 44611.5
563 | 266.7779 | 563 | 45365.9
563.2 | 230.5826 | 563.2 | 45925.2
563.4 | 438.6212 | 563.4 | 47422.1
563.6 | 199.973 | 563.6 | 46407.4
563.8 | 328.4162 | 563.8 | 47029.1
564 | 354.2922 | 564 | 47424.2
564.2 | 385.806 | 564.2 | 48242.1
564.4 | 263.0727 | 564.4 | 48164.3
564.6 | 342.9462 | 564.6 | 48649.4
564.8 | 318.9843 | 564.8 | 48786.8
565 | 314.3129 | 565 | 48204.3
565.2 | 288.6427 | 565.2 | 50876.8
565.4 | 214.5511 | 565.4 | 49284.9
565.6 | 271.5949 | 565.6 | 50449.8
565.8 | 265.1636 | 565.8 | 49603.2
566 | 315.1743 | 566 | 50904
566.2 | 259.0464 | 566.2 | 50757.3
566.4 | 141.1643 | 566.4 | 51327.2
566.6 | 284.6636 | 566.6 | 50779.9
566.8 | 342.4908 | 566.8 | 50627.4
567 | 330.4056 | 567 | 50826.7
567.2 | 217.9566 | 567.2 | 51813.9
567.4 | 305.9906 | 567.4 | 52087.1
567.6 | 274.6514 | 567.6 | 51374.2
567.8 | 369.8924 | 567.8 | 51555.8
568 | 364.2778 | 568 | 52253.6
568.2 | 409.2997 | 568.2 | 52664.6
568.4 | 358.7397 | 568.4 | 52529.9
568.6 | 219.7448 | 568.6 | 50939.7
568.8 | 372.3281 | 568.8 | 52050.5
569 | 162.8111 | 569 | 52050.4
569.2 | 360.7341 | 569.2 | 52142.4
569.4 | 207.8604 | 569.4 | 52837.9
569.6 | 303.526 | 569.6 | 52572.5
569.8 | 163.632 | 569.8 | 51084.3
570 | 380.9416 | 570 | 53083
570.2 | 253.6566 | 570.2 | 51618.4
570.4 | 202.3665 | 570.4 | 52116.4
570.6 | 389.7189 | 570.6 | 52531.9
570.8 | 196.9594 | 570.8 | 51522.1
571 | 460.9081 | 571 | 51238.9
571.2 | 242.7658 | 571.2 | 51674.5
571.4 | 320.2527 | 571.4 | 52660.6
571.6 | 332.8972 | 571.6 | 51481.1
571.8 | 346.2812 | 571.8 | 52640.3
572 | 333.1923 | 572 | 51383.7
572.2 | 269.3492 | 572.2 | 52552.8
572.4 | 308.2979 | 572.4 | 52174.9
572.6 | 276.0554 | 572.6 | 50433.8
572.8 | 498.3675 | 572.8 | 50495.9
573 | 206.0643 | 573 | 51780.9
573.2 | 232.1779 | 573.2 | 50420.3
573.4 | 436.4095 | 573.4 | 51069.8
573.6 | 285.7213 | 573.6 | 49819.7
573.8 | 311.913 | 573.8 | 50389.2
574 | 247.0796 | 574 | 50485.9
574.2 | 174.7367 | 574.2 | 49506.4
574.4 | 253.7587 | 574.4 | 50020.9
574.6 | 333.1026 | 574.6 | 48412.3
574.8 | 201.2437 | 574.8 | 49847.3
575 | 353.4412 | 575 | 47812.5
575.2 | 400.693 | 575.2 | 48993.2
575.4 | 294.6319 | 575.4 | 47559.2
575.6 | 341.4692 | 575.6 | 48161.8
575.8 | 256.0598 | 575.8 | 47650.5
576 | 156.857 | 576 | 47185.3
576.2 | 336.565 | 576.2 | 47453.7
576.4 | 270.9097 | 576.4 | 46946
576.6 | 331.5886 | 576.6 | 46735
576.8 | 144.1752 | 576.8 | 46244.3
577 | 365.7948 | 577 | 46385
577.2 | 299.6808 | 577.2 | 46627
577.4 | 353.7101 | 577.4 | 45453.5
577.6 | 394.1234 | 577.6 | 44941
577.8 | 306.6018 | 577.8 | 44414.5
578 | 300.6113 | 578 | 44933
578.2 | 246.7493 | 578.2 | 44741.7
578.4 | 450.6634 | 578.4 | 43576
578.6 | 463.3214 | 578.6 | 43717.5
578.8 | 268.2305 | 578.8 | 41958.7
579 | 282.8723 | 579 | 43475
579.2 | 316.8597 | 579.2 | 42548.1
579.4 | 119.6453 | 579.4 | 43003.2
579.6 | 372.3856 | 579.6 | 41816.8
579.8 | 345.5757 | 579.8 | 40774.5
580 | 312.2439 | 580 | 41860.4
580.2 | 188.4142 | 580.2 | 40879.7
580.4 | 359.8783 | 580.4 | 40892.5
580.6 | 374.6484 | 580.6 | 40504.3
580.8 | 299.1756 | 580.8 | 39839.1
581 | 210.5099 | 581 | 38712.2
581.2 | 445.0242 | 581.2 | 39447.7
581.4 | 238.5498 | 581.4 | 39660.8
581.6 | 376.5161 | 581.6 | 38258.2
581.8 | 349.2409 | 581.8 | 37993.3
582 | 163.4328 | 582 | 37676.7
582.2 | 295.0009 | 582.2 | 37350.4
582.4 | 281.9104 | 582.4 | 36512.3
582.6 | 205.438 | 582.6 | 35732
582.8 | 268.6755 | 582.8 | 36135.7
583 | 365.7209 | 583 | 36467.1
583.2 | 206.226 | 583.2 | 35445.9
583.4 | 388.5157 | 583.4 | 35777.9
583.6 | 179.135 | 583.6 | 33720.2
583.8 | 200.0728 | 583.8 | 33723.1
584 | 129.9003 | 584 | 34328.2
584.2 | 263.5965 | 584.2 | 33780.7
584.4 | 144.7659 | 584.4 | 33847
584.6 | 158.921 | 584.6 | 33269.4
584.8 | 131.4924 | 584.8 | 33075
585 | 208.8718 | 585 | 33273.8
585.2 | 287.5468 | 585.2 | 33267.2
585.4 | 429.5006 | 585.4 | 32301.9
585.6 | 245.4935 | 585.6 | 31299.9
585.8 | 189.2962 | 585.8 | 30974.9
586 | 217.9816 | 586 | 31932.2
586.2 | 274.7464 | 586.2 | 30574
586.4 | 182.5509 | 586.4 | 29713.9
586.6 | 418.7189 | 586.6 | 29689.9
586.8 | 183.0564 | 586.8 | 30313.4
587 | 133.0543 | 587 | 29351.6
587.2 | 348.6 | 587.2 | 28509.9
587.4 | 240.9462 | 587.4 | 28516.3
587.6 | 313.6095 | 587.6 | 29301.1
587.8 | 213.1209 | 587.8 | 27945.5
588 | 198.4588 | 588 | 27635.9
588.2 | 206.1028 | 588.2 | 27165.2
588.4 | 242.8435 | 588.4 | 27093.9
588.6 | 250.6413 | 588.6 | 26592.4
588.8 | 286.9913 | 588.8 | 26213.4
589 | 243.4611 | 589 | 26299.7
589.2 | 193.2957 | 589.2 | 26055.4
589.4 | 324.822 | 589.4 | 25970.2
589.6 | 222.6805 | 589.6 | 25492.6
589.8 | 274.4186 | 589.8 | 26261.6
590 | 274.114 | 590 | 25521.7
590.2 | 355.5025 | 590.2 | 24674.1
590.4 | 297.6353 | 590.4 | 24704.1
590.6 | 290.0367 | 590.6 | 25209.7
590.8 | 172.8352 | 590.8 | 23723.2
591 | 291.0317 | 591 | 24089.8
591.2 | 313.4072 | 591.2 | 23094.3
591.4 | 269.3424 | 591.4 | 24096.2
591.6 | 217.9075 | 591.6 | 23341.7
591.8 | 337.0525 | 591.8 | 22321.2
592 | 381.7978 | 592 | 22808
592.2 | 78.5276 | 592.2 | 22584.6
592.4 | 175.3019 | 592.4 | 22352.3
592.6 | 197.3781 | 592.6 | 21629.3
592.8 | 309.3048 | 592.8 | 22112
593 | 548.0654 | 593 | 21603.2
593.2 | 393.1044 | 593.2 | 21975.5
593.4 | 138.9485 | 593.4 | 21275
593.6 | 116.7987 | 593.6 | 21164.3
593.8 | 169.1822 | 593.8 | 20254.3
594 | 423.8799 | 594 | 20564.6
594.2 | 154.7188 | 594.2 | 20076.7
594.4 | 335.2765 | 594.4 | 19900.3
594.6 | 177.7732 | 594.6 | 19866.2
594.8 | 208.1072 | 594.8 | 18793.5
595 | 110.001 | 595 | 18880.8
595.2 | 269.0856 | 595.2 | 19714.1
595.4 | 201.2243 | 595.4 | 18341.8
595.6 | 254.3489 | 595.6 | 18526.8
595.8 | 224.7808 | 595.8 | 18497.8
596 | 179.5339 | 596 | 18574.4
596.2 | 340.3139 | 596.2 | 17653.7
596.4 | 179.5885 | 596.4 | 18728
596.6 | 149.866 | 596.6 | 18911.5
596.8 | 180.1117 | 596.8 | 18024.8
597 | 180.7219 | 597 | 17762.7
597.2 | 142.5858 | 597.2 | 17713.9
597.4 | 158.5125 | 597.4 | 17720.8
597.6 | 189.335 | 597.6 | 17499.7
597.8 | 390.1846 | 597.8 | 17157.3
598 | 282.3029 | 598 | 17246
598.2 | 182.7501 | 598.2 | 17481.6
598.4 | 275.2648 | 598.4 | 16690.2
598.6 | 236.9342 | 598.6 | 16936.2
598.8 | 159.6122 | 598.8 | 15494.8
599 | 277.204 | 599 | 16342.3
599.2 | 183.9847 | 599.2 | 15352.2
599.4 | 300.9189 | 599.4 | 16684.8
599.6 | 238.9006 | 599.6 | 16660.3
599.8 | 145.63 | 599.8 | 15349.5
600 | 240.2252 | 600 | 15236.7
600.2 | 185.5873 | 600.2 | 14892.4
600.4 | 162.0496 | 600.4 | 14428.4
600.6 | 359.0349 | 600.6 | 15485
600.8 | 312.3934 | 600.8 | 15112.8
601 | 328.7766 | 601 | 14399.4
601.2 | 60.1118 | 601.2 | 14583.7
601.4 | 68.1827 | 601.4 | 13854.3
601.6 | 418.3151 | 601.6 | 14233
601.8 | 140.2417 | 601.8 | 14990.2
602 | 4.7855 | 602 | 13986.6
602.2 | 164.4737 | 602.2 | 14659.3
602.4 | 227.9419 | 602.4 | 13859.8
602.6 | 300.1056 | 602.6 | 14227.2
602.8 | 132.9617 | 602.8 | 13298.9
603 | 237.6737 | 603 | 13592.8
603.2 | 262.2586 | 603.2 | 13396.4
603.4 | 230.8488 | 603.4 | 13470.5
603.6 | 109.6232 | 603.6 | 12961.2
603.8 | 190.8559 | 603.8 | 12883.2
604 | 287.799 | 604 | 12015.3
604.2 | 263.6228 | 604.2 | 12668.8
604.4 | 199.3271 | 604.4 | 13032.5
604.6 | 256.1778 | 604.6 | 12628.4
604.8 | 273.6184 | 604.8 | 12979.3
605 | 331.1418 | 605 | 12395.9
605.2 | 397.1908 | 605.2 | 12785.8
605.4 | 176.1614 | 605.4 | 12495.3
605.6 | 160.2018 | 605.6 | 12469.4
605.8 | 135.866 | 605.8 | 12085.9
606 | 276.2795 | 606 | 12343.4
606.2 | 218.4537 | 606.2 | 11717.6
606.4 | 268.4271 | 606.4 | 11410.7
606.6 | 433.1171 | 606.6 | 11992.5
606.8 | 219.2994 | 606.8 | 11027
607 | 261.5549 | 607 | 12321.6
607.2 | 352.8621 | 607.2 | 11169.4
607.4 | 71.2982 | 607.4 | 11159.1
607.6 | 79.901 | 607.6 | 11279.2
607.8 | 263.8795 | 607.8 | 11749.4
608 | 280.411 | 608 | 10613.6
608.2 | 390.2202 | 608.2 | 10872.9
608.4 | 372.9339 | 608.4 | 10450.8
608.6 | 306.4341 | 608.6 | 10552.4
608.8 | 315.0443 | 608.8 | 10477.3
609 | 181.1192 | 609 | 10957.7
609.2 | 198.6915 | 609.2 | 10377.3
609.4 | 299.9449 | 609.4 | 10553.2
609.6 | 309.5374 | 609.6 | 10246.1
609.8 | 149.0323 | 609.8 | 9749.8452
610 | 276.5985 | 610 | 9945.1078
610.2 | 259.5847 | 610.2 | 9832.4854
610.4 | 277.3111 | 610.4 | 9859.434
610.6 | 269.8294 | 610.6 | 10044.6
610.8 | 193.3175 | 610.8 | 9384.0883
611 | 330.619 | 611 | 9775.8732
611.2 | 313.8809 | 611.2 | 9619.7295
611.4 | 384.6143 | 611.4 | 9566.9734
611.6 | 229.7963 | 611.6 | 9815.8903
611.8 | 359.1427 | 611.8 | 9372.5103
612 | 212.6001 | 612 | 9045.1712
612.2 | 161.318 | 612.2 | 8757.3262
612.4 | 178.6801 | 612.4 | 7994.9807
612.6 | 57.224 | 612.6 | 8913.0787
612.8 | 222.5938 | 612.8 | 9110.6181
613 | 127.1031 | 613 | 9164.8813
613.2 | 57.6322 | 613.2 | 8031.3091
613.4 | 189.4177 | 613.4 | 7895.3331
613.6 | 189.4237 | 613.6 | 7987.1609
613.8 | 339.057 | 613.8 | 8268.3734
614 | 224.9669 | 614 | 8192.2804
614.2 | 146.6126 | 614.2 | 8361.7228
614.4 | 472.8525 | 614.4 | 8275.823
614.6 | 208.9273 | 614.6 | 7664.4482
614.8 | 23.0195 | 614.8 | 7367.1575
615 | 40.8385 | 615 | 7611.9875
615.2 | 49.8478 | 615.2 | 8271.2767
615.4 | 467.5663 | 615.4 | 7939.6281
615.6 | 183.9036 | 615.6 | 7551.6037
615.8 | 139.6173 | 615.8 | 7506.1775
616 | 103.9763 | 616 | 7814.6381
616.2 | 131.3447 | 616.2 | 7711.3423
616.4 | 293.443 | 616.4 | 7352.4287
616.6 | 213.0761 | 616.6 | 6986.6908
616.8 | 267.0433 | 616.8 | 7059.2144
617 | 195.0968 | 617 | 6937.264
617.2 | 167.8337 | 617.2 | 6693.1748
617.4 | 313.8128 | 617.4 | 7248.5716
617.6 | 195.9845 | 617.6 | 6353.6648
617.8 | 123.6383 | 617.8 | 6700.8442
618 | 224.3637 | 618 | 6684.1929
618.2 | 133.6303 | 618.2 | 6414.8599
618.4 | 188.705 | 618.4 | 6402.0906
618.6 | 152.4503 | 618.6 | 6672.2256
618.8 | 134.3216 | 618.8 | 5927.9565
619 | 291.1189 | 619 | 6244.1529
619.2 | 282.1179 | 619.2 | 6546.5571
619.4 | 199.9948 | 619.4 | 6290.7553
619.6 | 255.855 | 619.6 | 6065.3874
619.8 | 265.2207 | 619.8 | 6192.0057
620 | 396.0752 | 620 | 5879.6164
620.2 | 257.0576 | 620.2 | 5611.5725
620.4 | 285.8278 | 620.4 | 5451.7262
620.6 | 164.8503 | 620.6 | 5806.5667
620.8 | 90.0117 | 620.8 | 5351.3312
621 | 268.6205 | 621 | 5378.265
621.2 | 137.6393 | 621.2 | 5578.7074
621.4 | 184.6991 | 621.4 | 5541.8211
621.6 | 213.8586 | 621.6 | 5228.9278
621.8 | 166.7013 | 621.8 | 5392.7416
622 | 271.3078 | 622 | 5101.1604
622.2 | 319.7261 | 622.2 | 4647.412
622.4 | 300.7915 | 622.4 | 4886.2338
622.6 | 24.8675 | 622.6 | 4935.539
622.8 | 264.6612 | 622.8 | 4781.7062
623 | 332.4034 | 623 | 5204.1571
623.2 | 140.486 | 623.2 | 4719.8672
623.4 | 517.3411 | 623.4 | 4532.0093
623.6 | 131.5184 | 623.6 | 5086.8817
623.8 | -42.5736 | 623.8 | 4772.9121
624 | 102.8189 | 624 | 5057.1628
624.2 | 161.1612 | 624.2 | 4789.4875
624.4 | -91.5833 | 624.4 | 4238.4494
624.6 | 279.252 | 624.6 | 4348.5246
624.8 | 93.913 | 624.8 | 4566.5641
625 | 5.88 | 625 | 4201.0946
625.2 | 291.0422 | 625.2 | 4773.8462
625.4 | 252.2 | 625.4 | 3681.1718
625.6 | 104.4382 | 625.6 | 4101.136
625.8 | 114.6242 | 625.8 | 4145.4496
626 | 382.3445 | 626 | 4370.9046
626.2 | 254.3131 | 626.2 | 3945.9069
626.4 | 15.9196 | 626.4 | 4102.7695
626.6 | 55.8101 | 626.6 | 3798.1361
626.8 | 135.89 | 626.8 | 3630.3043
627 | 116.3074 | 627 | 3660.2658
627.2 | 327.3098 | 627.2 | 3974.6091
627.4 | 66.3276 | 627.4 | 3770.8857
627.6 | 308.1347 | 627.6 | 3668.5961
627.8 | 187.6788 | 627.8 | 3230.9695
628 | 97.2697 | 628 | 3556.0189
628.2 | 219.1588 | 628.2 | 3705.6679
628.4 | 97.501 | 628.4 | 3324.0699
628.6 | 158.8129 | 628.6 | 3114.8844
628.8 | 281.5745 | 628.8 | 3101.5232
629 | 159.5068 | 629 | 3119.6791
629.2 | 324.428 | 629.2 | 3202.5577
629.4 | 190.6762 | 629.4 | 3546.1973
629.6 | 314.8915 | 629.6 | 3225.0297
629.8 | 356.7089 | 629.8 | 3447.0949
630 | 171.5755 | 630 | 3035.3336
630.2 | 389.1854 | 630.2 | 3261.0198
630.4 | 151.6226 | 630.4 | 2848.3925
630.6 | 214.3185 | 630.6 | 3079.6885
630.8 | 162.634 | 630.8 | 3396.6534
631 | 383.0604 | 631 | 2916.529
631.2 | 236.9836 | 631.2 | 2678.1877
631.4 | 416.142 | 631.4 | 3097.2553
631.6 | 258.6374 | 631.6 | 3404.9473
631.8 | -109.8314 | 631.8 | 2921.5358
632 | 302.4884 | 632 | 2678.9456
632.2 | 207.8533 | 632.2 | 2520.4563
632.4 | 325.1484 | 632.4 | 2918.0563
632.6 | 336.9487 | 632.6 | 2308.9773
632.8 | 102.3637 | 632.8 | 2763.4961
633 | 166.5795 | 633 | 2149.5309
633.2 | 456.6812 | 633.2 | 2310.4896
633.4 | 92.627 | 633.4 | 2302.0004
633.6 | 362.8175 | 633.6 | 2256.082
633.8 | 266.0697 | 633.8 | 2671.4018
634 | 125.4401 | 634 | 2531.9254
634.2 | 485.0307 | 634.2 | 2315.9655
634.4 | -69.5948 | 634.4 | 2093.8108
634.6 | 453.6325 | 634.6 | 2580.9169
634.8 | 104.8814 | 634.8 | 2449.1856
635 | 214.3637 | 635 | 2188.0726
635.2 | 490.3613 | 635.2 | 2342.6647
635.4 | 215.2046 | 635.4 | 2346.2686
635.6 | 138.8261 | 635.6 | 2300.874
635.8 | 261.1134 | 635.8 | 2159.0556
636 | 205.9023 | 636 | 2186.7715
636.2 | 250.5807 | 636.2 | 1899.9742
636.4 | 195.6375 | 636.4 | 2075.7759
636.6 | 129.137 | 636.6 | 2223.1422
636.8 | 219.505 | 636.8 | 2495.9866
637 | 175.2484 | 637 | 1989.156
637.2 | 85.3421 | 637.2 | 2036.985
637.4 | 254.2339 | 637.4 | 1647.4588
637.6 | 322.5108 | 637.6 | 1750.3238
637.8 | 221.9442 | 637.8 | 1678.9349
638 | 346.6441 | 638 | 1662.0438
638.2 | 222.5691 | 638.2 | 1790.5199
638.4 | 200.1719 | 638.4 | 2025.2903
638.6 | 143.5184 | 638.6 | 2453.2901
638.8 | 224.0319 | 638.8 | 1782.8665
639 | 167.4769 | 639 | 1912.5969
639.2 | -50.591 | 639.2 | 1490.4744
639.4 | 329.0995 | 639.4 | 1654.3462
639.6 | 168.3179 | 639.6 | 1713.2839
639.8 | 157.1968 | 639.8 | 1486.7159
640 | 320.2209 | 640 | 1699.7214
640.2 | 251.2859 | 640.2 | 1509.8367
640.4 | 216.8127 | 640.4 | 1639.7264
640.6 | 53.7076 | 640.6 | 1632.9431
640.8 | 112.1865 | 640.8 | 1555.7587
641 | 371.1828 | 641 | 1642.0872
641.2 | 89.5137 | 641.2 | 1692.1075
641.4 | 384.9311 | 641.4 | 1545.8393
641.6 | 232.0724 | 641.6 | 1324.0041
641.8 | 316.066 | 641.8 | 1633.1893
642 | 90.4746 | 642 | 1257.5503
642.2 | 90.7924 | 642.2 | 1437.6198
642.4 | 150.7599 | 642.4 | 1133.7024
642.6 | 151.8469 | 642.6 | 1174.2858
642.8 | 211.9798 | 642.8 | 1367.0385
643 | 188.7906 | 643 | 1492.1654
643.2 | 116.313 | 643.2 | 1376.4752
643.4 | 79.9176 | 643.4 | 1549.7731
643.6 | 262.5643 | 643.6 | 1612.2818
643.8 | 372.8707 | 643.8 | 1397.6577
644 | 129.8324 | 644 | 1340.2251
644.2 | 386.7328 | 644.2 | 1171.2332
644.4 | 118.0675 | 644.4 | 1212.4669
644.6 | 327.4842 | 644.6 | 1558.746
644.8 | 315.247 | 644.8 | 1240.9373
645 | 118.7615 | 645 | 1405.8246
645.2 | 168.5702 | 645.2 | 1148.5051
645.4 | -104.5154 | 645.4 | 717.0308
645.6 | 206.507 | 645.6 | 1217.437
645.8 | 195.244 | 645.8 | 1069.7525
646 | 95.0148 | 646 | 1148.006
646.2 | 258.9122 | 646.2 | 809.161
646.4 | 196.0609 | 646.4 | 787.8922
646.6 | 323.5255 | 646.6 | 1196.1648
646.8 | 437.0246 | 646.8 | 1034.9874
647 | 71.0374 | 647 | 970.3627
647.2 | 351.2077 | 647.2 | 821.7171
647.4 | -43.3092 | 647.4 | 1078.9447
647.6 | 224.9554 | 647.6 | 850.8961
647.8 | 148.7499 | 647.8 | 723.8145
648 | 302.9231 | 648 | 572.3646
648.2 | 277.9522 | 648.2 | 997.4325
648.4 | 457.9566 | 648.4 | 883.974
648.6 | 408.4209 | 648.6 | 615.6137
648.8 | 266.6642 | 648.8 | 1069.4893
649 | 267.2307 | 649 | 722.689
649.2 | 33.8922 | 649.2 | 907.3705
649.4 | 125.3665 | 649.4 | 712.7782
649.6 | 465.7212 | 649.6 | 935.609
649.8 | 336.3456 | 649.8 | 937.7182
650 | 296.9986 | 650 | 838.2533
650.2 | 60.4502 | 650.2 | 665.8864
650.4 | 428.4017 | 650.4 | 734.29
650.6 | 325.2965 | 650.6 | 905.1654
650.8 | 298.5815 | 650.8 | 681.743
651 | 74.0785 | 651 | 884.6616
651.2 | 219.9757 | 651.2 | 789.8222
651.4 | 525.4429 | 651.4 | 751.0973
651.6 | 606.4834 | 651.6 | 847.6536
651.8 | 260.1504 | 651.8 | 686.5466
652 | 314.4431 | 652 | 847.9757
652.2 | 288.3737 | 652.2 | 543.0855
652.4 | 168.5748 | 652.4 | 945.9888
652.6 | 369.3828 | 652.6 | 1132.4101
652.8 | 222.7363 | 652.8 | 892.4784
653 | 250.6888 | 653 | 1014.7626
653.2 | 196.5127 | 653.2 | 547.8997
653.4 | 210.7427 | 653.4 | 736.3132
653.6 | 346.1681 | 653.6 | 844.6938
653.8 | 266.128 | 653.8 | 926.1971
654 | 157.1713 | 654 | 645.4581
654.2 | 89.7317 | 654.2 | 551.5463
654.4 | 185.3316 | 654.4 | 593.4295
654.6 | 280.8941 | 654.6 | 621.6197
654.8 | 267.4008 | 654.8 | 457.9324
655 | 323.1718 | 655 | 731.7862
655.2 | 145.7892 | 655.2 | 872.0103
655.4 | 159.2989 | 655.4 | 378.2744
655.6 | 91.0699 | 655.6 | 711.5405
655.8 | 132.423 | 655.8 | 366.8522
656 | 35.9472 | 656 | 698.4723
656.2 | 562.2793 | 656.2 | 741.1075
656.4 | 202.7409 | 656.4 | 714.5507
656.6 | 285.4548 | 656.6 | 535.4294
656.8 | 342.2611 | 656.8 | 716.0087
657 | 77.725 | 657 | 870.4354
657.2 | 469.4771 | 657.2 | 593.1768
657.4 | 190.3155 | 657.4 | 691.8818
657.6 | 442.0179 | 657.6 | 568.8804
657.8 | 275.2154 | 657.8 | 554.8448
658 | 149.066 | 658 | 976.3235
658.2 | 135.2943 | 658.2 | 360.6066
658.4 | 149.7149 | 658.4 | 627.8542
658.6 | 305.5844 | 658.6 | 516.6361
658.8 | 65.0842 | 658.8 | 644.0569
659 | 319.5364 | 659 | 404.3997
659.2 | 348.6773 | 659.2 | 575.1757
659.4 | 433.4473 | 659.4 | 974.0365
659.6 | 421.4209 | 659.6 | 891.7899
659.8 | 336.663 | 659.8 | 692.8503
660 | -34.3895 | 660 | 565.3389
660.2 | 208.5813 | 660.2 | 551.7474
660.4 | 238.3958 | 660.4 | 581.3849
660.6 | 80.4485 | 660.6 | 496.8308
660.8 | 253.6014 | 660.8 | 670.7286
661 | 182.0547 | 661 | 225.11
661.2 | 124.265 | 661.2 | -20.2054
661.4 | 356.254 | 661.4 | 546.1352
661.6 | 385.7591 | 661.6 | 255.5105
661.8 | -136.8909 | 661.8 | 444.4576
662 | 271.2683 | 662 | 241.3841
662.2 | 315.5922 | 662.2 | 504.3628
662.4 | 228.2705 | 662.4 | 462.236
662.6 | -35.1579 | 662.6 | 681.8209
662.8 | 361.1956 | 662.8 | 626.5795
663 | 140.9349 | 663 | 508.3381
663.2 | 303.6787 | 663.2 | 82.5346
663.4 | 171.1213 | 663.4 | 333.2271
663.6 | 364.8025 | 663.6 | 289.2029
663.8 | 216.9526 | 663.8 | 454.2248
664 | 454.8774 | 664 | 647.9043
664.2 | 113.4829 | 664.2 | 455.512
664.4 | 187.8676 | 664.4 | 768.5747
664.6 | 218.1558 | 664.6 | 411.7784
664.8 | 383.3697 | 664.8 | 580.343
665 | 368.3156 | 665 | 203.7607
665.2 | 309.173 | 665.2 | 563.653
665.4 | -21.0961 | 665.4 | 369.8598
665.6 | 401.5189 | 665.6 | 144.4315
665.8 | 266.2964 | 665.8 | 460.9247
666 | 281.7929 | 666 | 402.5506
666.2 | 130.223 | 666.2 | 432.9851
666.4 | 328.4533 | 666.4 | 327.8181
666.6 | 39.6503 | 666.6 | 450.3583
666.8 | 207.5181 | 666.8 | 207.337
667 | 100.825 | 667 | 712.6428
667.2 | 453.3651 | 667.2 | 345.7172
667.4 | 55.35 | 667.4 | 147.3702
667.6 | 302.1465 | 667.6 | 331.7905
667.8 | 86.3318 | 667.8 | 486.6586
668 | 256.109 | 668 | 364.2036
668.2 | 457.2695 | 668.2 | 333.6831
668.4 | 164.2006 | 668.4 | 287.8524
668.6 | 287.9187 | 668.6 | 102.3779
668.8 | 71.4625 | 668.8 | 273.2558
669 | 289.5148 | 669 | 678.7248
669.2 | 242.8354 | 669.2 | 367.1161
669.4 | 321.8445 | 669.4 | 274.974
669.6 | 338.5783 | 669.6 | 400.2896
669.8 | 480.663 | 669.8 | 480.8091
670 | 213.9228 | 670 | 779.1241
670.2 | -6.3174 | 670.2 | 388.0452
670.4 | 530.6449 | 670.4 | 294.1294
670.6 | 41.289 | 670.6 | 389.5455
670.8 | 294.9722 | 670.8 | 374.5642
671 | 533.9665 | 671 | 628.5786
671.2 | 121.0988 | 671.2 | 470.3273
671.4 | 296.6295 | 671.4 | 248.7295
671.6 | 73.7665 | 671.6 | 617.4477
671.8 | 491.2557 | 671.8 | 410.3152
672 | 315.2979 | 672 | 122.1981
672.2 | 251.3727 | 672.2 | 154.597
672.4 | -167.9642 | 672.4 | 447.0572
672.6 | 236.4225 | 672.6 | 139.4319
672.8 | 350.0256 | 672.8 | 351.5981
673 | -22.7936 | 673 | 481.3907
673.2 | 75.1539 | 673.2 | 368.6876
673.4 | 401.839 | 673.4 | 270.6871
673.6 | 337.2316 | 673.6 | 779.5003
673.8 | 223.4354 | 673.8 | 370.4384
674 | -55.9501 | 674 | 256.0703
674.2 | 108.7117 | 674.2 | 470.4771
674.4 | 207.8698 | 674.4 | 686.9512
674.6 | -6.612 | 674.6 | 175.2841
674.8 | 43.2505 | 674.8 | 374.5213
675 | -23.2908 | 675 | 192.9732
675.2 | 260.4979 | 675.2 | 493.6773
675.4 | 177.566 | 675.4 | 377.5639
675.6 | 866.5851 | 675.6 | 311.6146
675.8 | 26.88 | 675.8 | 144.3427
676 | 10.1171 | 676 | 649.2887
676.2 | 162.2971 | 676.2 | 295.8452
676.4 | 179.3426 | 676.4 | 43.883
676.6 | 264.0428 | 676.6 | 196.4214
676.8 | 129.019 | 676.8 | 282.2006
677 | 538.9404 | 677 | 265.635
677.2 | 454.5609 | 677.2 | 198.064
677.4 | 266.8561 | 677.4 | -58.1521
677.6 | 626.9823 | 677.6 | 370.0784
677.8 | 354.3082 | 677.8 | 526.2225
678 | 27.5645 | 678 | 217.5159
678.2 | 199.9736 | 678.2 | 681.687
678.4 | 304.4013 | 678.4 | 304.6324
678.6 | -93.6727 | 678.6 | 340.1263
678.8 | 426.9238 | 678.8 | 357.3294
679 | 706.6247 | 679 | 202.1303
679.2 | -24.4744 | 679.2 | 551.584
679.4 | 431.48 | 679.4 | 430.253
679.6 | 168.6335 | 679.6 | 115.6587
679.8 | 379.9606 | 679.8 | 503.2679
680 | -77.5818 | 680 | 169.2115
680.2 | 415.9918 | 680.2 | 646.0448
680.4 | 364.7764 | 680.4 | 382.4404
680.6 | -95.7426 | 680.6 | 578.9889
680.8 | 170.858 | 680.8 | 491.5712
681 | 28.5577 | 681 | 189.4246
681.2 | 332.8027 | 681.2 | 350.1901
681.4 | 351.8932 | 681.4 | 478.0599
681.6 | 82.6238 | 681.6 | -133.2956
681.8 | 263.5232 | 681.8 | 461.6118
682 | 100.7844 | 682 | 354.9737
682.2 | 210.0795 | 682.2 | 462.9734
682.4 | -25.4393 | 682.4 | 156.2939
682.6 | 794.3888 | 682.6 | 482.6869
682.8 | 522.6109 | 682.8 | 29.1892
683 | 120.5639 | 683 | 666.9122
683.2 | 415.3798 | 683.2 | 505.4115
683.4 | 139.9681 | 683.4 | 506.4355
683.6 | 195.7245 | 683.6 | 342.2555
683.8 | 48.2128 | 683.8 | 103.1897
684 | 251.8322 | 684 | 325.4183
684.2 | 159.8148 | 684.2 | 85.4823
684.4 | 364.8732 | 684.4 | 381.6174
684.6 | 403.32 | 684.6 | 85.7519
684.8 | 104.7045 | 684.8 | 122.8693
685 | -232.5877 | 685 | 553.4619
685.2 | 517.6164 | 685.2 | 535.2943
685.4 | 200.0969 | 685.4 | 424.9224
685.6 | 11.3303 | 685.6 | 180.9211
685.8 | 768.3617 | 685.8 | 502.6647
686 | 334.0621 | 686 | 125.3495
686.2 | 30.463 | 686.2 | 106.3294
686.4 | 374.0504 | 686.4 | 164.1304
686.6 | 757.682 | 686.6 | 393.1316
686.8 | -26.8006 | 686.8 | 299.3617
687 | 146.1445 | 687 | 607.7922
687.2 | 647.4941 | 687.2 | 473.9682
687.4 | -123.6844 | 687.4 | 223.8717
687.6 | 436.3792 | 687.6 | 341.5814
687.8 | 302.796 | 687.8 | -7.7475
688 | 303.8278 | 688 | 537.4189
688.2 | 381.9993 | 688.2 | -144.5862
688.4 | 148.6757 | 688.4 | 70.2359
688.6 | 559.4719 | 688.6 | 324.7161
688.8 | 854.488 | 688.8 | 129.683
689 | 228.5668 | 689 | -47.2484
689.2 | -165.8258 | 689.2 | 545.2494
689.4 | 785.274 | 689.4 | -166.4404
689.6 | 250.5465 | 689.6 | 270.462
689.8 | 508.6814 | 689.8 | 111.6084
690 | -107.3248 | 690 | 251.4423
690.2 | 252.0982 | 690.2 | 111.8652
690.4 | 533.349 | 690.4 | 112.3352
690.6 | 514.5133 | 690.6 | 333.4146
690.8 | 455.3428 | 690.8 | 293.7693
691 | 676.5096 | 691 | 496.09
691.2 | 154.2723 | 691.2 | 355.7326
691.4 | 700.8959 | 691.4 | 542.021
691.6 | 746.7326 | 691.6 | 337.8592
691.8 | 155.2666 | 691.8 | 278.2262
692 | 422.9404 | 692 | 155.3674
692.2 | 299.6261 | 692.2 | 155.8868
692.4 | 177.0947 | 692.4 | -8.2407
692.6 | 218.3921 | 692.6 | 301.1584
692.8 | 53.7072 | 692.8 | 425.8012
693 | 240.4669 | 693 | 219.3537
693.2 | 531.797 | 693.2 | 12.447
693.4 | -29.1933 | 693.4 | 554.0609
693.6 | 389.0896 | 693.6 | 137.6346
693.8 | -92.1671 | 693.8 | 138.1667
694 | 264.3288 | 694 | 12.5981
694.2 | 159.8141 | 694.2 | 12.6126
694.4 | 308.4041 | 694.4 | 286.6152
694.6 | 562.3338 | 694.6 | -50.6599
694.8 | 436.7975 | 694.8 | 329.6634
695 | 755.3607 | 695 | 118.5554
695.2 | 290.1331 | 695.2 | 352.6835
695.4 | 461.7013 | 695.4 | 374.8041
695.6 | 354.8572 | 695.6 | 226.398
695.8 | 657.551 | 695.8 | 55.813
696 | 12.924 | 696 | -115.9327
696.2 | 185.5881 | 696.2 | 271.4027
696.4 | 294.6126 | 696.4 | 293.8633
696.6 | 380.9081 | 696.6 | 1009.3445
696.8 | 318.6504 | 696.8 | 641.5765
697 | 711.8032 | 697 | 340.0996
697.2 | 624.7478 | 697.2 | -8.7421
697.4 | 647.7181 | 697.4 | -30.6983
697.6 | 275.6527 | 697.6 | 232.6735
697.8 | 474.0652 | 697.8 | 475.0371
698 | -52.9432 | 698 | 13.2549
698.2 | 234.2248 | 698.2 | 344.8653
698.4 | 146.3708 | 698.4 | 632.5907
698.6 | 435.1037 | 698.6 | 503.3437
698.8 | 303.1228 | 698.8 | 414.2659
699 | 259.4465 | 699 | 303.5897
699.2 | 439.3143 | 699.2 | 371.733
699.4 | 1134.5872 | 699.4 | -76.2371
699.6 | 58.6256 | 699.6 | 13.4859
699.8 | 487.466 | 699.8 | 396.441
700 | 13.588 | 700 | 397.3416
700.2 | 285.3353 | 700.2 | 58.8363
700.4 | 263.956 | 700.4 | 354.8
700.6 | 582.291 | 700.6 | 310.2471
700.8 | 515.5041 | 700.8 | 858.5309
701 | 127.9494 | 701 | 151.1808
701.2 | -147.1222 | 701.2 | 405.3553
701.4 | 313.8779 | 701.4 | -78.571
701.6 | 244.7641 | 701.6 | 152.435
701.8 | 408.3038 | 701.8 | 130.063
702 | 549.6411 | 702 | 572.4741
702.2 | -429.3095 | 702.2 | 762.7094
702.4 | 84.3195 | 702.4 | 317.4904
702.6 | 131.5287 | 702.6 | 459.8089
702.8 | -32.9059 | 702.8 | -56.5001
703 | -9.4102 | 703 | 84.9654
703.2 | 511.2767 | 703.2 | 298.4374
703.4 | 536.5614 | 703.4 | 61.7119
703.6 | 228.3024 | 703.6 | 466.5953
703.8 | 204.987 | 703.8 | 443.3273
704 | 14.3491 | 704 | 397.1736
704.2 | 302.3808 | 704.2 | 446.2193
704.4 | 449.2536 | 704.4 | -105.9923
704.6 | 473.7124 | 704.6 | -82.1809
704.8 | 38.8056 | 704.8 | 208.398
705 | 329.9316 | 705 | 305.846
705.2 | 379.4262 | 705.2 | 501.856
705.4 | -156.0146 | 705.4 | 355.9217
705.6 | 503.4579 | 705.6 | -9.7985
705.8 | 210.9979 | 705.8 | 358.8048
706 | -133.6403 | 706 | 88.7312
706.2 | 261.6779 | 706.2 | 336.5299
706.4 | 510.5271 | 706.4 | 510.4883
706.6 | 263.9319 | 706.6 | 511.5219
706.8 | 713.6891 | 706.8 | 439.0931
707 | 341.1607 | 707 | 414.5665
707.2 | -35.0002 | 707.2 | -35.1414
707.4 | 666.811 | 707.4 | -10.0413
707.6 | 65.4951 | 707.6 | 367.7383
707.8 | 722.4471 | 707.8 | 546.0185
708 | 724.8612 | 708 | 217.6592
708.2 | 193.1305 | 708.2 | 1005.1654
708.4 | 142.677 | 708.4 | 269.9543
708.6 | 168.9057 | 708.6 | 474.8322
708.8 | 553.2858 | 708.8 | -266.2451
709 | 143.6617 | 709 | 272.1789
709.2 | 272.8376 | 709.2 | 221.2664
709.4 | 685.5131 | 709.4 | 894.1572
709.6 | 715.1367 | 709.6 | 196.2279
709.8 | 171.0958 | 709.8 | -321.6358
710 | 119.6199 | 710 | 744.2887
710.2 | -244.8966 | 710.2 | 537.463
710.4 | 721.599 | 710.4 | 642.725
710.6 | -167.2946 | 710.6 | 487.8255
710.8 | 409.5264 | 710.8 | 672.3239
711 | 252.6433 | 711 | 173.7252
711.2 | 886.7494 | 711.2 | 993.1819
711.4 | -10.6109 | 711.4 | -382.0644
711.6 | 307.7029 | 711.6 | -10.6205
711.8 | -37.2465 | 711.8 | 122.7357
712 | 362.2862 | 712 | 495.6113
712.2 | 603.7408 | 712.2 | 390.5165
712.4 | 497.987 | 712.4 | 230.3482
712.6 | 579.1922 | 712.6 | 150.2636
712.8 | 123.9009 | 712.8 | -91.5302
713 | 124.4624 | 713 | 123.8725
713.2 | -200.6008 | 713.2 | 583.9588
713.4 | 422.5442 | 713.4 | 313.9612
713.6 | 587.6115 | 713.6 | 314.8865
713.8 | 561.6081 | 713.8 | 397.3832
714 | 235.0929 | 714 | 98.3736
714.2 | 235.263 | 714.2 | 262.0916
714.4 | 620.3415 | 714.4 | 428.3632
714.6 | 428.0348 | 714.6 | 566.6866
714.8 | 897.4307 | 714.8 | 209.2726
715 | 293.2197 | 715 | 568.9329
715.2 | 848.7207 | 715.2 | 708.3699
715.4 | 1127.8287 | 715.4 | 211.1126
715.6 | 294.9803 | 715.6 | 740.0649
715.8 | 463.8397 | 715.8 | 547.0839
716 | 297.2947 | 716 | 240.3236
716.2 | 1057.4851 | 716.2 | -95.2685
716.4 | 242.3851 | 716.4 | 16.8734
716.6 | 300.1385 | 716.6 | 214.0927
716.8 | 299.9943 | 716.8 | 639.1495
717 | 529.0772 | 717 | 472.835
717.2 | 216.0265 | 717.2 | 330.2632
717.4 | 45.7424 | 717.4 | 245.4077
717.6 | 503.2738 | 717.6 | 246.1514
717.8 | 304.7183 | 717.8 | -270.1603
718 | 74.831 | 718 | 190.2229
718.2 | 363.5401 | 718.2 | 103.8174
718.4 | -242.7094 | 718.4 | 480.6793
718.6 | 133.4621 | 718.6 | 654.76
718.8 | -244.9439 | 718.8 | 688.1232
719 | -11.6628 | 719 | -128.6275
719.2 | 46.8449 | 719.2 | 398.9403
719.4 | 486.0709 | 719.4 | 311.7305
719.6 | 578.1129 | 719.6 | 47.0638
719.8 | 76.6606 | 719.8 | -424.9048
720 | 491.0204 | 720 | 136.0349
720.2 | 254.7248 | 720.2 | 402.8665
720.4 | 702.3932 | 720.4 | 226.0483
720.6 | 316.6443 | 720.6 | 764.7551
720.8 | 77.6603 | 720.8 | 197.7229
721 | 347.9929 | 721 | 558.8807
721.2 | 198.1028 | 721.2 | 78.3585
721.4 | 440.5463 | 721.4 | 138.7917
721.6 | 684.0629 | 721.6 | 775.546
721.8 | 564.2861 | 721.8 | 472.1059
722 | 48.5213 | 722 | 899.0739
722.2 | 689.7521 | 722.2 | 506.5936
722.4 | 415.0401 | 722.4 | 1334.1613
722.6 | 662.462 | 722.6 | 600.5981
722.8 | 511.1212 | 722.8 | 388.2992
723 | -197.0089 | 723 | -43.0483
723.2 | 297.1131 | 723.2 | 482.8639
723.4 | 111.6374 | 723.4 | 546.3642
723.6 | 485.0824 | 723.6 | -168.1496
723.8 | 237.0104 | 723.8 | 143.5628
724 | 833.0545 | 724 | 18.7678
724.2 | 584.5782 | 724.2 | 582.9126
724.4 | 81.9952 | 724.4 | 333.7409
724.6 | 872.1553 | 724.6 | 616.725
724.8 | 462.0506 | 724.8 | -328.4071
725 | 398.8222 | 725 | 18.9915
725.2 | 559.2025 | 725.2 | 781.3751
725.4 | 528.5663 | 725.4 | 529.1087
725.6 | 754.2142 | 725.6 | 274.5174
725.8 | 338.9926 | 725.8 | 275.4391
726 | -44.926 | 726 | 532.9565
726.2 | 856.5433 | 726.2 | 19.3125
726.4 | -207.111 | 726.4 | 277.4713
726.6 | 19.4295 | 726.6 | 116.8698
726.8 | 768.2802 | 726.8 | 311.7968
727 | -110.8821 | 727 | -78.1509
727.2 | 936.6866 | 727.2 | 247.909
727.4 | 609.0241 | 727.4 | 314.8131
727.6 | -13.1398 | 727.6 | 315.2008
727.8 | 118.3273 | 727.8 | 877.7008
728 | 747.2751 | 728 | 151.9327
728.2 | 648.1244 | 728.2 | 417.999
728.4 | 186.0145 | 728.4 | 385.2279
728.6 | 486.6723 | 728.6 | 86.8295
728.8 | 120.258 | 728.8 | 1089.4167
729 | 589.1136 | 729 | 756.5904
729.2 | 389.6662 | 729.2 | 120.7473
729.4 | 323.1281 | 729.4 | 457.001
729.6 | 964.8663 | 729.6 | 87.5721
729.8 | 223.6814 | 729.8 | 54.0797
730 | 292.441 | 730 | 54.2301
730.2 | 531.6493 | 730.2 | 20.4139
730.4 | 598.6748 | 730.4 | -115.9253
730.6 | 532.538 | 730.6 | 431.2712
730.8 | 20.5646 | 730.8 | 20.5857
731 | 89.2099 | 731 | 1224.8774
731.2 | 675.0707 | 731.2 | 848.9895
731.4 | 886.2949 | 731.4 | -83.027
731.6 | -221.9856 | 731.6 | 990.5304
731.8 | 229.2684 | 731.8 | 646.3898
732 | 265.0059 | 732 | 403.9608
732.2 | 90.5893 | 732.2 | 440.5775
732.4 | 476.3784 | 732.4 | 476.4327
732.6 | 477.5736 | 732.6 | 898.2121
732.8 | 409.1688 | 732.8 | 1005.6389
733 | 1189.7637 | 733 | 339.1449
733.2 | 21.3092 | 733.2 | 56.7359
733.4 | 234.2598 | 733.4 | 56.8032
733.6 | -156.5655 | 733.6 | 698.3028
733.8 | 307.2942 | 733.8 | 842.2797
734 | 271.8377 | 734 | 631.6526
734.2 | -50.2879 | 734.2 | 21.5955
734.4 | 380.6988 | 734.4 | 814.5508
734.6 | 453.9084 | 734.6 | -14.4388
734.8 | 491.1251 | 734.8 | 94.0199
735 | 312.0517 | 735 | 94.1766
735.2 | 1329.0033 | 735.2 | 494.2685
735.4 | 167.6369 | 735.4 | 642.2223
735.6 | 277.9749 | 735.6 | 387.3486
735.8 | 644.872 | 735.8 | 498.6509
736 | 499.6257 | 736 | 905.4878
736.2 | 944.5613 | 736.2 | -125.1497
736.4 | -88.7918 | 736.4 | 317.8203
736.6 | 170.7078 | 736.6 | 96.1833
736.8 | -274.4985 | 736.8 | -424.1951
737 | 1141.2027 | 737 | 694.0945
737.2 | 619.7251 | 737.2 | 359.6513
737.4 | 584.6326 | 737.4 | 209.6052
737.6 | 547.9651 | 737.6 | 586.1422
737.8 | 1113.5501 | 737.8 | 737.2668
738 | -15.0655 | 738 | 361.8455
738.2 | 173.8152 | 738.2 | 779.3945
738.4 | 629.3733 | 738.4 | 667.0848
738.6 | 784.955 | 738.6 | 669.4444
738.8 | 557.9103 | 738.8 | 328.3947
739 | 635.1703 | 739 | 978.2759
739.2 | 789.825 | 739.2 | 444.9755
739.4 | 177.0552 | 739.4 | 483.7815
739.6 | 717.0365 | 739.6 | 138.7444
739.8 | 602.2753 | 739.8 | 796.4303
740 | 993.7609 | 740 | 1339.867
740.2 | 256.7443 | 740.2 | 915.0625
740.4 | 1272.424 | 740.4 | 101.2428
740.6 | 453.3681 | 740.6 | 414.0328
740.8 | 454.5572 | 740.8 | 376.1995
741 | 63.0894 | 741 | 731.0991
741.2 | 534.8837 | 741.2 | 574.2788
741.4 | 300.2877 | 741.4 | 220.853
741.6 | 656.6703 | 741.6 | 339.8928
741.8 | 341.7197 | 741.8 | 262.0107
742 | 382.1636 | 742 | 421.7317
742.2 | 663.8403 | 742.2 | 63.8289
742.4 | -16.022 | 742.4 | 143.3358
742.6 | 544.1971 | 742.6 | 623.4692
742.8 | -136.5786 | 742.8 | 786.9766
743 | -16.1594 | 743 | 505.9936
743.2 | 508.5915 | 743.2 | 386.8371
743.4 | 550.7261 | 743.4 | 347.5004
743.6 | 714.9984 | 743.6 | 632.6176
743.8 | 24.4059 | 743.8 | 268.6383
744 | 228.2149 | 744 | 555.8401
744.2 | 65.5299 | 744.2 | 1128.801
744.4 | 189.1662 | 744.4 | 353.5372
744.6 | 106.8007 | 744.6 | 312.6379
744.8 | 684.962 | 744.8 | 354.0243
745 | 1343.7579 | 745 | 520.789
745.2 | 273.4707 | 745.2 | 521.9019
745.4 | 232.4503 | 745.4 | 770.808
745.6 | 857.4199 | 745.6 | 66.4619
745.8 | -183.6229 | 745.8 | 1318.5459
746 | 946.0015 | 746 | 569.1888
746.2 | 403.6648 | 746.2 | 696.652
746.4 | 869.5933 | 746.4 | -185.2726
746.6 | 279.1339 | 746.6 | -396.1074
746.8 | 406.4372 | 746.8 | 194.6033
747 | 621.642 | 747 | 407.246
747.2 | 835.9627 | 747.2 | 663.7415
747.4 | 1092.2797 | 747.4 | 793.73
747.6 | 1223.9356 | 747.6 | 711.0172
747.8 | 839.2465 | 747.8 | 111.5611
748 | 111.9785 | 748 | 799.8006
748.2 | 285.1956 | 748.2 | 1147.8535
748.4 | -407.2216 | 748.4 | 242.3697
748.6 | 113.0764 | 748.6 | 69.3926
748.8 | 810.9159 | 748.8 | 808.8229
749 | -104.9622 | 749 | 1116.5362
749.2 | 946.2076 | 749.2 | 550.9487
749.4 | 421.8322 | 749.4 | 421.1592
749.6 | 599.1814 | 749.6 | 70.3661
749.8 | 865.3833 | 749.8 | 424.2645
750 | 1706.4376 | 750 | -636.3327

## References
[1] Simple PL QY = (Em(sample) - Em(solvent)) / (Scatter(solvent) - Scatter(sample))
[2] Corrected PL QY from Edinburgh Instruments PL FLS980 Integrating Sphere Manual
[3] Porres, L et al, J. Fluoresc. 2006, 16 (2), 267-273; also de Mello, Wittman and Friend, Adv Mater, 1997, 9, 3
Jacobian correction: Phys. Chem. Lett. 2013, 4, 19, 3316–3318
2 NDF meas
Input %T